- OBJECTIVE:

The objective of this SOP is to describe the procedure for the calibration of digital pH meter to ensure accurate pH measurements by standardizing the pH meter using certified buffer solutions

- SCOPE:

To provide written procedure for calibration of Digital pH Meter in the Chemical Laboratory.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- QC Analyst shall be responsible for operation, calibration and maintenance of the instrument and its related documentation as per the SOP.

- QC Head shall be responsible to provide training to the concerned analyst before the implementation of the SOP and to ensure the procedure is followed as per the SOP.

- ABBREVIATION AND DEFINITIONS

SOP – standard Operating Procedure

- Materials

- pH meter

- pH electrode

- Temperature probe

- pH 4.00, 7.00 buffer solutions

- WFI water

- Clean beakers

- Lint-free tissue.

- PROCEDURES

CALIBRATION:

- Ensure that the electrode of the pH meter is connected properly to the display unit and the unit is plugged to the power supply.

- Switch on the power supply to the unit.

- Rinses the electrode with WFI water and wipe it with tissue paper gently.

- Set the “TEMPERATURE C” control to the temperature of the selected buffer solution.( 25Oc + 2oC)

- Immerse the electrode and temperature probe into PH 7.00 buffer. Set the function selector switch to pH position and adjust “CALIBRATE “control until the meter displays the precise pH value of the buffer solution.

- Press the “Cal” button to enter calibration mode.

- Wait for the reading to stabilize (5-10secs), then press enter to confirm.

- Now set the function selector switch to STAND BY.

- Remove the electrode from the buffer solution. Wash with distilled or deionized water.

- Dip the combined electrode into another buffer solution of 4.01 for second point (Set 4 pH Value).

- Set the “TEMPERATURE C” control to the buffer solution temperature.

- Set the function selector switch to pH position. Adjust the “SLOPE” correction control, until the meter displays the pH value of the selected buffer solution. Wait for stabilization and confirm calibration as in step 3. (For some meters, you might press CAL again or a sequence of buttons). Check that the correct readings are obtained with both solutions without further adjustment.

- Remove the electrode from the buffer solution. Wash with Water for Injection.

- If the readings are not in the stipulated limits, repeat the calibration again.

- If still no good results, inform the QC Manager about it. Do not use the pH meter when the calibration is not good.

- If the result of calibration is ok, then go for measuring the pH of the sample.

- The pH meter should be calibrated at least once a day

- Acceptable slope for electrode is 90-105%, readings within +0.05 pH of buffer value

- Record calibration details (date, time, buffer values, temperature, slope) in a log book

- Store the electrode in an appropriate solution (like saturated KCl or pH 7 buffer with KCl) after rinsing. Avoid storing in distilled or deionized water, which can cause damage.

- REFERENCE AND ENCLOSURE

- Operation Manual supplied by the manufacturer.

- SOP for maintenance of Laboratory Instrument/equipment.

- USP General Chapter 791

- Instrument /equipment/ equipment usage logbook.

- Enclosure 1- Digital pH Meter Calibration Log - BML-QDG-QA-036-F01

- Enclosure 2: pH buffer Preparation log sheet- BML-QDG-QA-036-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | Quality Control Analyst |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |