- OBJECTIVE:

To establish a standard procedure for the calibration and verification of the temperature accuracy of the hot air oven used in the Microbiology Laboratory.

- SCOPE:

This SOP applies to all ovens used for drying, sterilization, or other specified purposes in the Microbiology laboratories.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Laboratory Assistants/Analysts: Performing internal calibrations, recording data, and maintaining logs.

- Quality Control Head: Overall accountability.

- External Agency: Performing annual, external calibration.

- ABBREVIATION AND DEFINITIONS

Nil

- PROCEDURES

- Ensure the oven is clean, free from residues, and not in use.

- Obtain the calibrated standard thermometer and verify its certificate.

- Place the oven on a stable, level surface.

- Set the oven to specified temperatures (e.g., 100°C and 150°C, or as per requirement).

- Allow the oven to reach and stabilize at the set temperature (usually 30-60 min).

- Place the standard thermometer on the upper and lower shelves (or multiple points for accuracy).

- Keep the thermometer in the oven for 15-30 minutes to equilibrate.

- Note the temperature reading on the standard thermometer and compare it to the oven's display.

- Record observations in the Oven Calibration Record according to the below template

| Set temperature | Thermometer | Oven reading | Tolerance (±2°C) | Thermometer reading | Remark |
| --- | --- | --- | --- | --- | --- |

- Repeat readings at different shelf locations and set points.

- Ensure the observed temperature is within the acceptable tolerance (e.g., ±2.0°C) of the set temperature.

Internal calibration: Quarterly

External Calibration: Annually by a qualified external agency.

- Record all observations, deviations, and adjustments in the designated logbook.

- Apply a calibration status sticker (Pass/Fail) to the oven.

Precautions:

- Do not open the oven door unnecessarily during calibration.

- Use tongs and gloves when handling hot items.

- REFERENCE AND ENCLOSURE

- Training on SOP

- Manual supplied by the supplier

7.0 TRAINING INVOLVEMENT

| Quality Control Head | Quality Control Analysts |
| --- | --- |
| QA Doc. Unit staff | Laboratory Assistants |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |