| S/N | TEST | DESCRIPTIONS OF PROCEDURE |
| --- | --- | --- |
|  | DESCRIPTION | Describe the colour and other characteristics of the material. |
|  | SOLUBILITY | Determine solubility of material in water and methylene chloride. |
|  | BOILING POINT | Determine the Boiling Point (About 780C) |
|  | IDENTIFICATION | First identification A, B Second identification A, C, D A. Determine the relative density of the material. B. Infrared absorption spectrophotometry, comparing the spectrum obtained with that of Ph. Eur. Reference spectrum ethanol (96%). C. Mix 0.1 mL with 1 mL of a 10 g/L solution of potassium permanganate R and 0.2 mL of dilute sulphuric acid R in a test tube. Cover immediately with a filter paper moistened with a freshly prepared solution containing 0.1 g of sodium nitroprusside R and 0.5 g of piperazine hydrate R in 5 mL of water R. Note the coloration formed on the paper after a few minutes and what happen after 10 – 15 minutes. D. To 0.5 mL add 5 mL of water R, 2 mL of dilute sodium hydroxide solution R, then slowly add 2 mL of 0.05 M iodine. Note any precipitate formed within 30 minutes. |
|  | APPEARANCE | Observe the clarity and colour of the material when compared with water. Dilute 1.0 mL to 20 mL with water R, After standing for 5 minutes, observe and record the clarity of the dilution in comparison with water. |
|  | ACIDITY OR ALKALINITY | To 20 mL, add 20 mL of carbon dioxide - free water and 0.1 mL of phenolphthalein solution R. Observe the colour of the solution. Add 1.0 mL of 0.01 M sodium hydroxide. Note any coloration of the solution. |
|  | RELATI VE DENSITY | Determine the relative density of the material. |
|  | ABSORBANCE | Measure the absorbance at 240 nm, between 250 nm and 260 nm and also between 270 nm – 340 nm. Examine the absorbance between 235 nm and 340 nm in a 5 cm cell using water as the compensation liquid. Note any smoothness in the absorption curve. |
|  | VOLATILE IMPURITIES | Gas chromatography. Test solution (a) The substance to be examined. Test solution (b) Add 150 µL of 4-methylpentan-2-ol R to 500.0 mL of the substance to be examined. Reference solution (a) Dilute 100 µL of anhydrous methanol R to 50.0 mL with the substance to be examined. Dilute 5.0 mL of the solution to 50.0 mL with the substance to be examined. Reference solution (b) Dilute 50 µL of anhydrous methanol R and 50 µL of acetaldehyde R to 50.0 mL with the substance to be examined. Dilute 100 µL of the solution to 10.0 mL with the substance to be examined. Reference solution (c) Dilute 150 µL of acetal R to 50.0 mL with the substance to be examined. Dilute 100 µL of the solution to 10.0 mL with the substance to be examined. Reference solution (d) Dilute 100 µL of benzene R to 100.0 mL with the substance to be examined. Dilute 100 µL of the solution to 50.0 mL with the substance to be examined. Column: -- material: fused silica, -- size: l = 30 m, Ø = 0.32 mm, -- stationary phase: Poly[(cyanopropyl)(phenyl)][dimethyl]siloxane R (film thickness 1.8 µm) carrier gas helium for chromatography R Linear velocity 35 cm/s Split ratio 1: 20 Temperature: Time (min); Temp.(0C). column; 0 – 12 12 --32 32 -- 42; 40 40 – 240 240. Injection port; 200. detector; 280. Detection Flame ionization Injection 1 µL System suitability Reference solution (b): --- resolution: minimum 1.5 between the first peak (acetaldehyde) and the second peak (methanol). Limits : -- methanol: in the chromatogram obtained with test solution(a): not more than half the area of the corresponding peak in the chromatogram obtained with reference solution (a) (200 ppm V/V) -- acetaldehyde + acetal: maximum of 10 ppm V/V, expressed as acetaldehyde. Calculate the sum of the contents of acetaldehyde and acetal in parts per million (V/V) using the following expression: 10 x AE + 30 x CE x 44.05 AT -- AE CT -- CE 118.2 AE =area of the peak due to acetaldehyde in the chromatogram obtained with test solution (a), AT = area of the peak due to acetaldehyde in the chromatogram obtained with reference solution (b), CE =area of the peak due to acetal in the chromatogram obtained with test solution (a), CT =area of the peak due to acetal in the chromatogram obtained with test solution (c), 44.05 = molecular mass of acetaldehyde, 118.2 = molecular mass of acetal, -- benzene: maximum 2 ppm V/V Calculate the content of benzene in parts per million (V/V) using the following expression: 2BE BT -- BE BE = area of the peak due to benzene in the chromatogram obtained with the test solution (a). BT = area of the peak due to benzene in the chromatogram obtained with the test solution (d). If necessary the identity of benzene can be confirmed using another suitable chromatographic system (stationary phase with a different polarity). -- total of other impurities: in the chromatogram obtained with test solution (b): not more than the area of the peak due to 4-methylpentan-2-ol in the chromatogram obtained with test solution (b) (300 ppm), -- disregard limit: 0.03 times the area of the peak due to 4-methylpentan-2-ol in the chromatogram obtained with test solution (b) (9ppm). |
|  | RESIDUE ON EVAPORATION | Evaporate 100 mL to dryness on a water – bath and dry at 100 – 105 0C for I hour. The residue weighs a maximum of 2.5 mg. |
|  | ETHANOL | Calculated from the relative density value using the alcoholimetric tables. Apparent Density = 997.2 x Specific Gravity |
|  | MICROBIAL ENUMERATION TEST | Determination of total aerobic microbial count (TAMC) and total combined yeasts and molds count (TYMC) by the pour plate method. - Slightly open the lid of the petri plate and dispense 1 ml of diluted sample in the center of the Petri plate using a sterile calibrated pipette. - Open the lid of the prepared media. (soya bean digest agar for TAMC and Sabouraud Dextrose Agar for TYMC) - Pour about 15-20ml of the sterilized molten media at the appropriate temperature above the sample. - Close the lid of the plate then mix the sample in the media properly by gently swirling the plate. The plate is generally swirled in an “S” or “8” shape. - Allow the media to completely solidify. - Simultaneously perform the negative control by plating 1ml of diluent used instead of the sample for each medium. - Incubate the plate in an inverted position under suitable incubation conditions (For TAMC Incubate the plates of Soybean Casein Digest Agar at 30°C-35°C for 48hours and For TYMC incubate the plates of Sabouraud Dextrose Agar at 20°C - 25°C for 72hours Test for Absence of Pathogen Test for E. coli: Using same dilution - Transfer 1ml of the sample into a sterile McCartney bottle containing 9ml of sterile MacConkey broth. - Incubate the MacConkey broth bottle at 300C-350C for 24 hours. - After incubation shake the tube and transfer loop full portion of MacConkey broth on the surface of the pre incubated Eosin Methylene Blue Agar petri plate. - Incubate plates at 30°C to 35°C for 48 hours. - For negative control, proceeding same as above except without sample, streak a portion of incubated sterile MacConkey broth medium on the surface of pre incubated Eosin Methylene Blue Agar. - Observe and record the presence or absence of specified microorganisms. - Include positive and negative controls in each batch of tests. - Verify the performance of media and reagents used in the analysis. Interpretation of Results - The bacterial count should not exceed 102cfu/ml while that of the fungal count should not be more than 102cfu/ml. There should be absence of fecal coliform (E coli). - If growth occurs in the negative control plate, the test is invalid, failed negative control must be investigated - The total aerobic microbial count (TAMC) is considered to be equal to the number of CFU found using casein soya bean digest agar; if colonies of fungi are detected on this medium, they are counted as part of TAMC. - The total combined yeasts/mould count (TYMC) is considered to be equal to the number of CFU found using Sabouraud dextrose agar; if colonies of bacteria are detected on this medium, they are counted as part of TYMC - When an acceptance criterion for microbiological quality is prescribed it is interpreted as follows: 101 CFU : maximum acceptable count＝20 102 CFU: maximum acceptable count＝200, 103 CFU: maximum acceptable count＝2000. |

| - CURRENT REVISION | - SUPERSEDE | - REASON FOR REVIEW |
| --- | --- | --- |
| - 00 | - NIL | - FIRST ISSUE |
| - 01 | - 00 | - Review of the procedures - Change in header and footer format |