Autoclave Validation Protocol

|  | WRITTEN BY: | REVIEWED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

|  | APRROVEDBY: | AUTHORIZED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

1.0 Table of Contents

| S./N | TITLE | PAGE NO. |
| --- | --- | --- |
| 1. | APPROVAL PAGE | 1 |
| 2. | TABLE OF CONTENTS | 2 |
| 3. | OBJECTIVE, SCOPE, EQUIPMENT DESCRIPTION, VALIDATION TEAM | 3 |
| 4. | RESPONSIBILITIES, DEFINITIONS | 4 |
| 5. | LOADED CHAMBER HEAT PENETRATION STUDY | 5 |
| 6. | FULL LOAD, PARTIAL LOAD | 5 |
| 7. | IDENTIFICATION OF CRITICAL CONTROL | 6 |
| 8. | METHOD HEAT DISTRIBUTION STUDY | 7 |
| 9. | HEA T PENETRATION STUDY | 8-9 |
| 10. | SCHEMATIC DIAGRAM OF DELAMA AUTOCLAVE (SENSORS LOCATION) | 10 |
| 11. | SCHEMATIC DIAGRAM OF PHARMALAB AUTOCLAVE (SENSORS LOCATION) | 11 |
| 12. | FULL LOADING WITH PRODUCTS, BIOLOGICAL CHALLENGE TEST, TEST CYCLE | 12 |
| 13. | VALIDATION FREQUENCY, REQUALIFICATION | 13 |
| 14. | PROBE REQUIREMENT, DOCUMENTATION | 14 |
| 15. | EQUIPMENT , PLACEMENT OF PROBES, RUNNING STUDY, EVALUATION DATA | 15 |
| 16. | ACCEPTANCE CRITERIA | 16 |
| 17. | VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE ( CHECKLIST REPORT) | 17-18 |
| 18. | VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE ( HEAT DISTRIBUTION STUDY) | 18-20 |
| 19. | VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE ( HEAT PENETRATION STUDY) | 20-21 |
| 20. | VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE ( PERFORMANCE) | 22-23 |
| 21. | CERTIFICATION | 24 |

- OBJECTIVE:

- The purpose of this document is to provide the documented evidences which prove that, the performance of Sterilizer is in accordance to the specifications and meet the requirements of the product, process, safety, regulatory bodies and cGMP obligations.

- To ensure that the current loading pattern have uniform heating pattern.

- To ensure that heat distribution, heat penetration carried out for the three consecutive cycles complies with our in-house requirements.

- To ensure a Sterility Acceptance Limit (SAL) value greater than 10-6 is achieved with the process parameter

- Scope: The scope of this validation exercise is verification and compilation of the study conducted and documentation on the validation process.

- EQUIPMENT DESCRIPTION:

| - DELAMA AUTOCLAVE- Industrial double door autoclave | - PHARMALAB AUTOCLAVE: Industrial double door autoclave |
| --- | --- |
| Equipment Identification Number: BL/AC/PM/001 | Equipment Identification Number: BL/AC/PM/002 |
| Location: Production floor | Location: Production floor |
| Capacity : 3000 pouches | Capacity: 6000pouches |
| Purpose To sterilize the filled pouches |  |

- VALIDATION TEAM:

- Autoclave Operator.

- Production Supervisor

- Microbiologist.

- Supervisor Engineering Department.

- Quality Control Head

- Chief Engineer

- Quality Assurance Manager

- RESPONSIBILITIES:

| Autoclave Operator | Machine operation and execution of the protocol. |
| --- | --- |
| Production supervisor | Supervision of the operators and equipment |
| Supervisor Engineering Department. | To conduct heat distribution and penetration studies. |
| Microbiologist. | Preparation of the protocol To conduct the Biological indicator test/ chemical Strip. Documentation of Results |
| Chief Engineer | Preparation of the protocol Utility support |
| Quality Control Head | Execution of the protocol. To verify and document the results gathered during the validation of Sterilizer |
| QA Manager | Approval |
| Supt. Pharm./ Factory Manager | Authorization |

- DEFINITIONS

7.1 Hot water shower process:

Heating, sterilizing and cooling is done by circulating demineralized water which is pumped from the chamber via a heat exchanger and distributed over the product. During the whole process support pressure is maintained in the chamber to prevent damages of the product.

7.2 Empty chamber study:

To understand the geometric heat distribution study, empty chamber study is done (3 times) without product load with all empty trays.

This test establishes the uniformity of temperature distribution in the superheated sterilizer chamber and identifies the location of the coolest point in the chamber.

7.3 LOADED CHAMBER HEAT DISTRIBUTION STUDIES

This test establishes the heat distribution with load in sterilizer, and as well establishes the slowest to heat point for when loaded at the established sterilization phase cycle time.

Probes located on the trolleys are situated on the trays perpendicular to the length of trolleys.  None of the exposed probe tips should touch product or metal surface.

| prerequisite | - Pre-calibrated temperature data loggers with RTD (Resistance Temperature Detector) probes |
| --- | --- |
| Acceptance criteria | - Locate the cool-spot in the chamber for a complete cycle. - This shall be established for both a partial and full load. |

7.4 LOADED CHAMBER HEAT PENETRATION STUDY

This test aims to verify the effectiveness of the sterilization cycle for each load configuration. Test should be conducted in triplicate and should be able to demonstrate that each cycle is capable of providing the required lethality and must demonstrate reproducibility for loading pattern, product sizes and also the nature of the solution to be sterilized

| prerequisite | - Pre-calibrated temperature data loggers with RTD (Resistance Temperature Detector) probes |  |  |
| --- | --- | --- | --- |
| Acceptance criteria | - The Fo value after three runs must be lesser than the sterilization cycle run. - Identification of cold spot - Average Temperature throughout the chamber does not fluctuate by more than 3°C. (For cycles with sterilization of 116°C. All temperatures are within 116°C to 118°C |  |  |

7.5 Full Load:

The maximum permissible quantity that can be loaded in the sterilizer or maximum permissible yield of the batch, not exceeding the earlier

7.6 Partial Load:

A sterilizer load containing one (1) trolley with one (1) full layer or more units missing of product.

In Delama Autoclave, For 1000ml NLT 584 units ,500ml NLT 1017 units & 100ml NLT 3667 units

In Pharmalab Autoclave, For 1000ml NLT 875units , 500ml, NLT1525 units & 100ml NMT 5500units

The partial loaded trolleys must be completed with empty trays. When sterilization of a partial load is employed, a chamber qualified sterilization full load may qualified for the minimum partial load by performing three (3) combination heat penetration / distribution / biological indicator study.

7.7 Minimum Load

The minimum quantity of the sterilization lot, the minimum quantity placed on the trays shall be placed at the middle of the sterilizer during qualification study.

- IDENTIFICATION OF CRITICAL CONTROL MONITORING PARAMETER:

Before proceeding for validation of Autoclave, following parameters to be checked:

- Biological Indicator spore concentration must be 106.

- Used media must be tested for its growth promoting test.

- Status of Pressure Gauge

- Temperature Gauge display on the Autoclave

- Status of Sensors

- Status of data logger with all probes before and after the validation using standard thermometer

- Time

- Heat penetration

- Cold spot identification

Acceptance criteria:

- Since Overkill approach will be used. It is assumed that the product to be sterilized is having bio load of 1 million CFU. i.e. Log of 1,000,000 + 6 log (required to achieve 12 log reduction).

We have 6 log + 6 log =12 log reduction

- To determine the cycle time, it is important to note that Geobacillus thermophilus spore D value is between 1 -3 minutes. To determine the cycle time the company make use of 3.33 minutes to give extra level of assurance for the terminal sterilization

- Cycle time: 12*3.33= 40Minutes Approximately

- Method

9.1 Heat Distribution and Heat Penetration Study

9.2 Materials and instruments used:

- Mercury thermometer

- Chemical test strips

- Biological test

9.3 HEAT DISTRIBUTION STUDIES (EMPTY CHAMBER):- To verify the temperature uniformity throughout the chamber and to locate the cold spot in Empty Chamber.
Procedure:- Insert 18 no of temperature sensors inside the chamber through the validation port of sterilizer. Seal the port with silicon sealant to ensure that no steam leakage during operations of sterilizer. Fix all the probe at different location in the sterilizer so that sensors do not touch the metallic surface of the chamber. Connect the temperature sensors to the data logger, which can print the actual temperature and pressure at different locations.
After completion of sterilization cycle check the thermograph in the data logger for attaining set temperature and pressure during the sterilizing hold period. If any deviation observed repeat the cycle after taking necessary correction.
Acceptance criteria:-Temperature distribution within the chamber must be between 116°C to 118°C at all location during the sterilization period (dwell time).
9.4 HEAT DISTRIBUTION STUDIES (LOADED CHAMBER):- To verify the temperature uniformity throughout the chamber and to locate the cold spot in loaded Chamber.
Procedure:- Insert 18 no of temperature sensors inside the chamber through the validation port of sterilizer. Seal the port with silicon sealant to ensure that no steam leakage during operations of sterilizer. Fix all the probe at different location in the sterilizer so that sensors do not touch the metallic surface of the chamber. Load the article as per loading pattern in the autoclave chamber. Loaded chamber heat distribution study shall be performed separately for all loading patterns. One cycle shall be performed for each loading type for loaded chamber heat distribution study. Connect the temperature sensors to the data logger, which can scan and print the actual temperature and pressure at different locations.
After completion of sterilization cycle check the thermograph in the data logger for attaining set temperature and pressure during the sterilizing hold period. If any deviation observed repeat the cycle after taking necessary correction.
Acceptance criteria:-Temperature distribution within the chamber must be between 116°C to 118°C at all location during the sterilization period (dwell time).
9.5 HEAT PENETRATION STUDIES:– In order to verify sterilizing temperature has been reached in each load subjected to moist heat sterilization, it is necessary to conduct heat penetration studies. This study is conducted to ensure that the coolest unit within a pre-defined loading pattern (including minimum and maximum loads of different characteristics) will consistently be exposed to sufficient heat lethality (minimum “F0”).
Procedure:- Insert 18 no of temperature sensors inside the chamber through the validation port of sterilizer. Seal the port with silicon sealant to ensure that no steam leakage during operations of sterilizer. Fix the probes at different location inside the pouches to be used for the study in the sterilizer so that sensors do not touch the metallic surface of the chamber. It is important to note that bracketing approach will be used for this study, i.e Product with highest viscosity for each batch size of 100ml, 500ml and 1000ml shall be picked for the study. Loading pattern in the autoclave chamber to be attached to the study report and this will vary according to the autoclave and also the batch sizes. Loaded chamber heat distribution study shall be performed separately for all loading patterns. Arrange the load as specified keep at least 18 biological indicators and 18 thermo chemical indicators shall be used for each cycle. Load placed at the identified cold points must have indicator and temperature sensor in all three runs. One biological indicator and thermo chemical indicator along with external temperature sensor shall be placed at drain point in all three cycles. Perform the sterilization by operating the program specified for each load type as per standard operating procedure and start the data logger and steam sterilizer simultaneously. After the sterilization cycle is completed, stop data logger and open the sterilizer and take out the biological indicator and thermo chemical indicator from load and send to microbiological lab for testing. The biological indicator shall aseptically inoculate into sterile soybean casein digest media (SCDM) and incubated for 7 days -14 days and check the thermo chemical indicator for the compliance as per manufacturer recommendation for colour change (i.e. Brown). Take out temperature chart/data logger and inbuilt temperature recorder, report of biological indicator. Check against acceptance criteria for compliance and determine the Fo value and compare against acceptance criteria. Take out external temperature sensor from the chamber and perform vacuum leak rate test.
Acceptance criteria:-Temperature distribution within the chamber must be between 116°C to 119°C at all location during the sterilization period (dwell time).
Sterilization temperature should be maintained for NLT 40 minute for minimum 18 thermo couple during hold time. Biological indicator (Geobacillus stearothermophillus) should show complete sterilization (i.e. no growth after incubation).

- SCHEMATIC  DIAGRAM OF DELAMA AUTOCLAVE  WITH  LOCATION OF TEMPERATURE SENSORS (INSIDE VIEW) HEAT MAPPING STUDIES

16

13

10

7

4

1

16

13

Trolley 3

Trolley 2

Trolley 1

up

10

7

11

4

1

5

2

6

3

17

14

11

8

5

2

8

17

14

Middle

6

3

18

15

12

9

18

15

12

9

Down

Cold spot identification

Cold spot identification

- SCHEMATIC  DIAGRAM OF PHARMALAB  AUTOCLAVE  WITH  LOCATION OF TEMPERATURE SENSORS (INSIDE VIEW) FOR HEAT MAPPING STUDIES

Trolley 4

18

1

14

10

16

6

2

Trolley 3

Trolley 2

Trolley 1

up

15

7

3

11

17

Middle

4

16

12

8

Down

17

5

9

13

Cold spot identification

Cold spot identification

Controllers and timers were set as per sterilization timings and temperature requirement of the sterilizer:

  - Full load with products

This involves load of the Pharmalab sterilizer with maximum load of 6000 pouches of bags filled with products and De- lama Autoclave loaded with 3000 pouches of bags filled with products using the acceptable tray load pattern. Three consecutive batches are used to get sufficient data for evaluation.

The following parameters were set and observed.

PARAMETERS (operational Temperature, operational pressure, Time)

- Sterilization pressure at 1.5 Bar

- Sterilization temperature not less than 116oC

- Hold time  40mins

- Test also involves the use of Thermalog Chemical Strip. This indicates complete sterilization with change in colour and movement of the chemical indicator from reject to the Accept zone.

- BIOLOGICAL CHALLENGE TEST

- This test outlines the destruction of a resistant microbial challenge by a moist heat cycle, at the sterilization cycle time established by the Fo studies.

- The microbiological (EZTEST) test kit containing spores of Geobacillus stearothermophilus were also used. This involves the placement of the test tubes at the worst case zone (Near the autoclave doors, the middle and the bottom part in the chamber). Then, the kits will be introduced into a prepared Tryptone soya broth and Fluid thioglycolate medium, incubated at 600c for 7-14 days after sterilization, and thereby should be no colour changes. Also the sterility test is done on autoclaved raw water sample.

- TEST CYCLE

| S/No | TEST PARAMETER | CYCLES |
| --- | --- | --- |
| 1 | EMPTY | 3 |
| 2 | HEAT DISTRIBUTION- EMPTY LOAD | 3 |
| 3 | HEAT DISTRIBUTION- MAXIMUM LOAD | 3 |
| 4 | HEAT PENETRATION- MAXIMUM LOAD | 3 |

- VALIDATION FREQUENCY

| S/NO | PARAMETER | STANDARD | ACCEPTANCE CRITERIA | FREQUENCY |
| --- | --- | --- | --- | --- |
| 1 | Heat distribution | Pre-calibrated thermocouple | ± 3.0°C | Yearly |
| 2 | Heat penetration | Pre-calibrated thermocouple | ± 3.0°C | Yearly |
| 3 | Physical indicator | Chemical strip | Colour change | Every cycle |
| 4 | Biological challenge Test | Geobacillus stearothermophilus | No turbidity | Biannually |

- REQUALIFICATION

- Sterilization cycle:

For requalification, one study is required. Use only full loads for requalification. Requalification of partial loads is not required unless Chamber modifications are made.

Each Chamber (or) process must be re-qualified every 180days  15days from the date of most recent qualification (or) requalification tests.  If it exceeds six (6) months for a process the Chamber will not be used for sterilization process.

Note: The Chamber qualification/validation interval will be 180days  15days.  If the re-qualification of the Chamber is performed within  15days of the due date, the next due date is calculated from the date the re-qualification/ Validation was due rather than from the date performed.

Product related:

- Initial qualification are required whenever trolley loading is changed, load distribution (or) maximum loading. (Note: Minor alterations as a result of trolley structure maintenance do not require qualification).

- A point qualification pattern is required.

Following any maintenance other than stated above, the system and a consideration must be verified against the appropriate System Requirement Documents for conformance

Other changes requiring qualification:

- Testing in the form of qualification study shall be performed when the process parameters (or) process events may potentially be impacted by a control system (or) software change example of these parameters are events include but are not limited to

a) Temperature.

b) Pressure

c) Cooling

d) Water levels

e) Alarms

f)  Heat gain changes

- Probe requirements:

- At least 18 probes must be functional besides the recorder probe.

- Recorder probe: Probes are to be located at the recorder for Pharmalab Autoclave and Delama Autoclave qualification and requalification/validation studies.

The thermocouple measuring junction must not contact any metal surface.

- DOCUMENTATION

To be considered an acceptable study, records of all qualification studies must include, at minimum, the following properly completed and reviewed documents.

- Validator data print out.

- Thermocouple calibration records

- Manually recorded data.

- Drawings indicating Probe placements.

-Drawings indicating tray arrangement

- Copy of sterilizer temperature and pressure chart and the operator's record for the load involved.

-Microbiology study data (if required).

- A Qualification Study Summary.

- A record of the qualification status of the sterilizer should be maintained and available to all sterilizer operators and personnel releasing routine cycles.

- EQUIPMENT

- Calibrated thermocouple probes.

- Product load: The load used for the study contains pouches filled with water for injection. QA must be present at the loading and unloading of probes and spores units.

- Spore units must be marked / labeled for easy identification. In routine production loads chemical indicators are used instead of Biological Indicators.

- PLACEMENT OF PROBES FOR PENETRATION STUDY:

- A system is used to describe the location of units on sterilizer trolleys.

- Place 18 probes in various trolleys and various rows, so that all trolleys and rows are covered during the Heat Penetration study.  The fill volume must be within the specified range for the product being tested.

- If a microbiological study is required, place a spore unit in appropriate place in the chamber

- RUNNING THE STUDY

- The position of probes during Heat Distribution study with load and without load studies should remain same.

- Place penetration / distribution probes on the trolleys and load into the Chamber. Probes must be placed in the same channel used for calibration.

- Set controllers and timers as per Sterilization timings and temperatures requirement of the Autoclave

- Record probe temperature at 2 Minutes intervals.

- Record the elapsed time since cycle start.

- The study may be aborted once if a cycle aberration occurs..

- EVALUATION OF DATA

Evaluate the following data from the Process Data printout to determine the acceptability of the study:

- Distribution Studies: Determine the following information. Include the exposed probe in the calculation. Do not include the first minute during exposure.

- Minimum temperature during exposure period.

- Maximum temperature during exposure period.

- Maximum temperature range (Δ: T) during exposure period. This is determined by calculating the difference between high and low probes at each minute. The largest difference is the maximum delta T.

- Penetration Studies: F0  at exposure end and cycle end.

- ACCEPTANCE CRITERIA:

- All thermocouples used in sterilizer qualification study must be calibrated before the start of study and meet all acceptance criteria (0.3C for the sterilizer temperature probes) If it does not meet the acceptance criteria the calibration will be repeated.

- Recalibration/Verification of thermocouples is done after the study. (If it does not meet the acceptance criteria  the study is On Hold.) A maximum of 5 studies may be run before recalibration/Verification.

- The cycles in which studies were done must meet all cycle operation requirements with no aberrancies.

- Chamber Temperature Recorder Chart vs. Exposed Probe: During the timed exposure period (except for the first minute after exposure start), the probe by the recorder must be within 0.5 °C of the Chamber temperature recorder chart.

- There must be at least 8 functional distribution probes, 8 functional penetration probes.

- Spore population study on Biological indicators must meet the requirements.

- Microbiological tests on sterilizer water shall meet the requirements.(<10CFU/1000ml)

- Check the gauge and recorder calibration accuracy. If a deviation is found, apply this factor to the readings found during the study. The study may be accepted if the corrected Chamber pressure per gauge and pressure recorder chart is within limits and within ± 0.15 bar of each other.

- Maximum delta T in exposure may not exceed 2.5C

- The F0  value of each functional penetration probe must be calculated

- VALIDATION OF DELAMA AND PHARMALAB AUTOCLAVE (CHECKLISTS REPORT)

| S. No. | Critical parameters to be met | Acceptance Criteria | Observation |
| --- | --- | --- | --- |
|  | Verify all gauges & instruments are ok Temp. Chart / Recorder. Temp. Controller. Pressure Gauges. | All records must be within the acceptable limits |  |
|  | Empty Chamber Heat Distribution Cycle. (No. of cycles: 3) Set temperature: 116 - 118oC. Sterilization time: 40mins. Pressure bar: 1.5 – 1.7 | Throughout the Sterilization phase all temperatures must be within 3oC range |  |
|  | Machine parts: - Safety valves for chamber, streamline ,air inlet - Pressure gauge for chamber, streamline, air inlet - Air pressure regulator - Non return valves - Heat Exchanger - Door gasket - Pressure switch/ pressure transmitter - Temperature sensor - Chart recorder - Programmable logic control (PLC) working | Must conform with set parameters |  |
|  | Autoclaving with pouches Set temperature: 116 - 118oC. Sterilization time: 40 mins. Biological indicators should be used in all 3 Cycles. Chemical indicators should be used in all 3 Cycles. | Throughout the sterilization phase all temperatures are within 2oC range. The colour of the Biological indicator with samples taken from the identified cold spot should not change in the inoculation broth after the incubation period of 7 to 14days at 45oC. Initial White colour of chemical indicator placed at the identified cold spot should change to black and cross from reject zone to accept zone of the strip |  |
|  | Cold spot identification During sterilization cycle, arranging the probes across the loaded chamber has the goal of assuring sterilization is attained across the load | Although the cold spot might be lagging before reaching the sterile hold temperature of 116°C. It must not drop below 116°C sterile hold temperature throughout the sterile hold time of 40mins. |  |

- VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE (HEAT DISTRIBUTION STUDY)

| S/N | PARAMETERS | ACCEPTANCE CRITERIA | HEAT DISTRIBUTION STUDY (NO LOAD) |  |  | HEAT PENETRATION STUDY (FULL LOAD) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 |  | Trial 3 |  |
| 1. | Thermalog Chemical strips - Point 1 - Point 2 - Point 3 - Point 4 - Point 5 - Point 6 - Point 7 - Point 8 - Point 9 - Point 10 - Point 11 - Point 12 - Point 13 - Point 14 - Point 15 - Point 16 - Point 17 - Point 18 | Chemical Strips must change colour and move from unsafe to safe zone |  |  |  |  |  |  |  |  |
| 2. | Microbiological Test Kits - Point 1 - Point 2 - Point 3 - Point 4 - Point 5 - Point 6 - Point 7 - Point 8 - Point 9 - Point 10 - Point 11 - Point 12 - Point 13 - Point 14 - Point 15 - Point 16 - Point 17 - Point 18 | Test kit must be sterile after autoclave |  |  |  |  |  |  |  |  |
| 3. | Heating | Average Temperature difference between probe and thermometer must not be more than ± 2oC |  |  |  |  |  |  |  |  |
| 4. | Sterility test( sampled from cold spot) | Must be sterile |  |  |  |  |  |  |  |  |
| 5. | Bacteria Endotoxin | ≤ 0.25IU/ml or less than 1.150c sum temperature rise for three rabbits. |  |  |  |  |  |  |  |  |

- VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE ( HEAT PENETRATION STUDY) SEPARATE FOR EACH BATCH SIZE OF 100ML, 500ML & 1000ML.

| S/N | PARAMETERS | ACCEPTANCE CRITERIA | HEAT PENETRATION STUDY (FULL LOAD) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Trial 1 |  | Trial 2 |  | Trial 3 |  |  |
| 1. | Thermalog Chemical strips - Point 1 - Point 2 - Point 3 - Point 4 - Point 5 - Point 6 - Point 7 - Point 8 - Point 9 - Point 10 - Point 11 - Point 12 - Point 13 - Point 14 - Point 15 - Point 16 - Point 17 - Point 18 | Chemical Strips must change colour and move from unsafe to safe zone |  |  |  |  |  |  |  |
| 2. | Microbiological Test Kits - Point 1 - Point 2 - Point 3 - Point 4 - Point 5 - Point 6 - Point 7 - Point 8 - Point 9 - Point 10 - Point 11 - Point 12 - Point 13 - Point 14 - Point 15 - Point 16 - Point 17 - Point 18 | Test kit must be sterile after autoclave |  |  |  |  |  |  |  |
| 3. | Heating | Average Temperature difference between probe and thermometer must not be more than ± 2oC |  |  |  |  |  |  |  |
| 4. | Sterility test( sampled from cold spot) | Must be sterile |  |  |  |  |  |  |  |
| 5. | Bacteria Endotoxin | ≤ 0.25IU/ml or less than 1.150c sum temperature rise for three rabbits. |  |  |  |  |  |  |  |

- VALIDATION REPORT OF DELAMA AND PHARMALAB AUTOCLAVE (PERFORMANCE)

FIRST FIRSTBATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

SECOND BATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

THIRD BATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

____________________	_______________                   _________________________

Technical Officer	                   Production Pharmacist                             Quality Assurance Officer

Remark:    _________________________________________________________________________________

Certification:

__________________                    __________________________              _____________________

Engineering Manager                            Production Manager                                        QA Manager

_________________________

Factory Manager/ Supt. Pharmacist