Autoclave Validation Format Approval

|  | WRITTEN BY: | REVIEWED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: | SENIOR ANALYST | QC HEAD |
| Date: |  |  |

|  | APRROVED BY: | AUTHORIZED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: | QA MANAGER | SUPT. PHARMACIST |
| Date: |  |  |

1.0 Table of Contents

| S./N | TITLE | PAGE NO. |
| --- | --- | --- |
| 1. | APPROVAL PAGE | 1 |
| 2. | TABLE OF CONTENTS | 2 |
| 3. | EQUIPMENT DESCRIPTION | 3 |
| 4. | VALIDATION TEAM | 3 |
| 5. | RESPONSIBILITIES | 3 |
| 6. | SCHEMATIC DIAGRAM OF DELAMA AUTOCLAVE ( TEMP. ARRANGEMENT OF SENSORS) | 4 |
| 7. | SCHEMATIC DIAGRAM OF PHARMALAB AUTOCLAVE ( TEMP. ARRANGEMENT OF SENSORS) | 5 |
| 8. | VALIDATION REPORT OF PHARMALAB/DELAMA AUTOCLAVE ( CHECKLISTS REPORT) | 6-7 |
| 9. | VALIDATION REPORT OF PHARMALAB/DELAMA AUTOCLAVE ( HEAT DISTRIBUTION STUDY) | 8-9 |
| 10. | VALIDATION REPORT OF PHARMALAB/DELAMA AUTOCLAVE ( HEAT PENETRATION STUDY) | 10-11 |
| 11. | VALIDATION REPORT OF PHARMALAB /DELAMA AUTOCLAVE ( PERFORMANCE) | 12 |
| 12. | CERTIFICATION | 13 |

EQUIPMENT DESCRIPTION:

| - DELAMA AUTOCLAVE- Industrial double door autoclave | - PHARMALAB AUTOCLAVE: Industrial double door autoclave |
| --- | --- |
| Equipment Identification Number: BL/AC/PM/001 | Equipment Identification Number: BL/AC/PM/002 |
| Location: Production floor | Location: Production floor |
| Purpose To sterilize the filled pouches |  |

- VALIDATION TEAM:

- Autoclave Operator.

- Production Supervisor

- Microbiologist.

- Chemist

- Technical Officer

- Quality Control Head

- Chief Engineering

- Quality Assurance Manager

- RESPONSIBILITIES:

| Autoclave Operator | Machine operation and execution of the protocol. |
| --- | --- |
| Production supervisor | Supervision of the operators and equipment |
| Supervisor Engineering Department. | To conduct heat distribution and penetration studies. |
| Microbiologist. | Preparation of the protocol To conduct the Biological indicator test/ chemical Strip. Documentation of Results |
| Quality Control Head | Execution of the protocol. To verify and document the results gathered during the validation of Sterilizer |
| Chief Engineer | Utility support |
| QA Manager | Approval |
| Supt. Pharm./ Factory Manager | Authorization |

- SCHEMATIC DIAGRAM OF DELAMA AUTOCLAVE  WITH LOCATION OF TEMPERATURE SENSORS (INSIDE VIEW)

16

13

10

7

4

1

16

13

Trolley 3

Trolley 2

Trolley 1

up

10

7

11

4

1

5

2

6

3

17

14

11

8

5

2

8

17

14

Middle

6

3

18

15

12

9

18

15

12

9

Down

Cold spot identification

Cold spot identification

- SCHEMATIC DIAGRAM OF PHARMALAB  AUTOCLAVE  WITH LOCATION OF TEMPERATURE SENSORS (INSIDE VIEW)

Trolley 4

18

1

14

10

16

6

2

Trolley 3

Trolley 2

Trolley 1

up

15

7

3

11

17

Middle

4

16

12

8

Down

17

5

9

13

Cold spot identification

Cold spot identification

VALIDATION OF AUTOCLAVE (CHECKLISTS REPORT)        DELAMA AUTOCLAVE

PHARMALAB AUTOCLAVE

| S. No. | Critical parameters to be met | Acceptance Criteria | Observation |
| --- | --- | --- | --- |
|  | Verify all gauges & instruments are ok Temp. Chart / Recorder. Temp. Controller. Pressure Gauges. | All records must be within the acceptable limits |  |
|  | Empty Chamber Heat Distribution Cycle. (No. of cycles: 3) Set temperature: 116 - 118oC. Sterilization time: 40mins. Pressure bar: 1.5 – 1.7 | Throughout the Sterilization phase all temperatures must be within 3oC range |  |
|  | Machine parts: - Safety valves for chamber, streamline ,air inlet - Pressure gauge for chamber, streamline, air inlet - Air pressure regulator - Non return valves - Heat Exchanger - Door gasket - Pressure switch/ pressure transmitter - Temperature sensor - Chart recorder - Programmable logic control (PLC) working | Must conform with set parameters |  |
|  | Autoclaving with pouches Set temperature: 116 - 118oC. Sterilization time: 40 mins. Biological indicators should be used in all 3 Cycles. Chemical indicators should be used in all 3 Cycles. | Throughout the sterilization phase all temperatures are within 2oC range. The colour of the Biological indicator with samples taken from the identified cold spot should not change in the inoculation broth after the incubation period of 7 to 14days at 45oC. Initial White colour of chemical indicator placed at the identified cold spot should change to black and cross from reject zone to accept zone of the strip |  |
|  | Cold spot identification During sterilization cycle, arranging the probes across the loaded chamber has the goal of assuring sterilization is attained across the load | Although the cold spot might be lagging before reaching the sterile hold temperature of 116°C. It must not drop below 116°C sterile hold temperature throughout the sterile hold time of 40mins. |  |

- VALIDATION REPORT          PHARMALAB AUTOCLAVE (HEAT DISTRIBUTION STUDY)

DELAMA AUTOCLAVE           (HEAT DISTRIBUTION STUDY)

| S/N | PARAMETERS | ACCEPTANCE CRITERIA | HEAT DISTRIBUTION STUDY (NO LOAD) |  |  | HEAT PENETRATION STUDY (FULL LOAD) |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 |  | Trial 3 |  |
| 1. | Thermalog Chemical strips Point 1 Point 2 Point 3 Point 4 Point 5 Point 6 Point 7 Point 8 Point 9 Point 10 Point 11 Point 12 Point 13 Point 14 Point 15 Point 16 Point 17 Point 18 | Chemical Strips must change colour and move from unsafe to safe zone |  |  |  |  |  |  |  |  |
| 2. | Microbiological Test Kits Point 1 Point 2 Point 3 Point 4 Point 5 Point 6 Point 7 Point 8 Point 9 Point 10 Point 11 Point 12 Point 13 Point 14 Point 15 Point 16 Point 17 Point 18 | Test kit must be sterile after autoclave |  |  |  |  |  |  |  |  |
| 3. | Heating | Average Temperature difference between probe and thermometer must not be more than ± 2oC |  |  |  |  |  |  |  |  |
| 4. | Sterility test( sampled from cold spot) | Must be sterile |  |  |  |  |  |  |  |  |
| 5. | Bacteria Endotoxin | ≤ 0.25IU/ml or less than 1.150c sum temperature rise for three rabbits. |  |  |  |  |  |  |  |  |

VALIDATION REPORT             PHARMALAB AUTOCLAVE ( HEAT PENETRATION STUDY)

DELAMA AUTOCLAVE          (HEAT PENETRATION STUDY)

100ML                 500ML       &            1000ML.

| S/N | PARAMETERS | ACCEPTANCE CRITERIA | HEAT PENETRATION STUDY (FULL LOAD) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Trial 1 | Trial 2 | Trial 3 |  |
| 1. | Thermalog Chemical strips Point 1 Point 2 Point 3 Point 4 Point 5 Point 6 Point 7 Point 8 Point 9 Point 10 Point 11 Point 12 Point 13 Point 14 Point 15 Point 16 Point 17 Point 18 | Chemical Strips must change colour and move from unsafe to safe zone |  |  |  |  |
| 2. | Microbiological Test Kits Point 1 Point 2 Point 3 Point 4 Point 5 Point 6 Point 7 Point 8 Point 9 Point 10 Point 11 Point 12 Point 13 Point 14 Point 15 Point 16 Point 17 Point 18 | Test kit must be sterile after autoclave |  |  |  |  |
| 3. | Heating | Average Temperature difference between probe and thermometer must not be more than ± 2oC |  |  |  |  |
| 4. | Sterility test( sampled from cold spot) | Must be sterile |  |  |  |  |
| 5. | Bacteria Endotoxin | ≤ 0.25IU/ml or less than 1.150c sum temperature rise for three rabbits. |  |  |  |  |

- VALIDATION REPORT    PHARMALAB AUTOCLAVE (PERFORMANCE)

DELAMA AUTOCLAVE        (PERFORMANCE)

FIRST FIRSTBATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

SECOND BATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

THIRD BATCH

DATE: _____________________________________________

PRODUCT NAME: ____________________________________

BATCH NUMBER: ____________________________________

| S/N | PARAMETERS | SPECIFICATION | RESULT |
| --- | --- | --- | --- |
| 1. | Temperature |  |  |
| 2. | Hold time |  |  |
| 3. | Pressure |  |  |

____________________	_______________                   _________________________

Technical Officer	                   Production Pharmacist                             Quality Assurance Officer

Remark:    _________________________________________________________________________________

Certification:

__________________                    __________________________              _____________________

Engineering Manager                            Production Manager                                        QA Manager

_________________________

Factory Manager/ Supt. Pharmacist