VALIDATION MASTER PLAN

| Document no. | : | BML/QA/VMP/002 |
| --- | --- | --- |
| Supersedes no. | : | BML/QA/VMP/001 |
| Effective Date | : | SEPTEMBER 2022 |
| Next Review | : | AUGUST 2025 |
| Page nos. | : | 1 to 39 |

| VALIDATION MASTER PLAN |  |  | DATE: Sept, 2022 |  |
| --- | --- | --- | --- | --- |
| INDEX |  |  |  |  |
|  |  | Title |  | Page |
|  | 1.0 | VALIDATION MASTER PLAN Approval |  | 3 |
|  | 2.0 | Introduction |  | 4 |
|  | 3.0 | Concept of qualification |  | 7 |
|  | 4.0 | Revalidation |  | 16 |
|  | 5.0 | Facility description |  | 16 |
|  | 6.0 | Description of building and equipment |  | 17 |
|  | 7.0 | Validation program and overview |  | 19 |
|  | 8.0 | Calibration program summary |  | 24 |
|  | 9.0 | Preventive maintenance program summary |  | 25 |
|  | 10.0 | Key standard operating procedures (sops) |  | 25 |
|  | 11.0 | Validation of building |  | 25 |
|  | 12.0 | Validation of utility system |  | 26 |
|  | 13.0 | Process flow, control variables, and measured responses. |  | 29 |
|  | 14.0 | Qualification of laboratory equipment |  | 30 |

| VALIDATION MASTER PLAN |  | DATE: Sept, 2022 |
| --- | --- | --- |
| 1.0 | VALIDATION MASTER PLAN Approval |  |
|  | The following are responsible for Approval of VALIDATION MASTER PLAN S/N; Activity; Designation/Department; Sign; Date. 1; Prepared by:. 2; Reviewed by:. 3; Approved by:. 4; Authorised by:. |

| VALIDATION MASTER PLAN |  |  | Sept, 2022 |  |
| --- | --- | --- | --- | --- |
| 2.0 | INTRODUCTION - INTRODUCTION - BIOMEDICAL LIMITED is an indigenous Pharmaceutical Manufacturing Company specialized in the production of intravenous solutions in flexible PVC containers under Bioflex brand. Other formulations include Oral liquid preparations - The facility was incorporated in 1978, the company commenced commercial production on 21st August 1981 to meet the national and international regulatory and cGMP requirements. - The manufacturing is supported by Engineering, Production, Quality control and Quality Assurance department. All these departments are covered under the validation activities. - The validation Master plan is a dynamic document which provides complete over-view of validation program. - It describes the overall objective, intention approach for establishing performance adequacy of equipment, utilities, processes and system. - It identifies the scope of validation, application validation protocol-reports, procedures and frequency of validations. |  |  |  |
| 2.1 | BUIDING DESCRIPTION The design and material used for the construction of building is according to cGMP regulations. All walls, ceilings, and flooring inside the production area are designed as per cGMP regulations. The building is divided into two, - Infusion Plant - Production area - Service floor for utilities. - QA/AC - Admin. - Oral Liquid plant (a)Production area (b)Service floor for utilities. ( c) QA/QC |  |  |  |
| 2.2 | Project description:- |  |  |  |
|  | This validation master plan (VMP) specifies and coordinates all qualification/ validation activities to ensure the production of finished dosage meet up to accepted international standards. It also specifies the responsibilities for various validation procedures and helps to plan the necessary activities. The production of each product is divided into the following processing stages: - Receipt of raw materials and sampling. - Interim storage of raw materials. - Weighing before usage. - Manufacturing. - Packaging. - Storage of finished product - Dispatch |  |  |  |
| 2.3 | Principle:- |  |  |  |
|  | Validation in general requires a detailed preparation and careful planning of the various steps in the production process of products manufactured in BIOMEDICAL Limited, Ilorin (Nigeria) in addition, all work involved shall be carried out in a structured way according to formally authorised working and administrative procedures. Validation is documented evidence that provides a high degree of assurance that a specific process will consistently produce a product that meets its predetermined specifications and quality attributes. - Multidisciplinary approach: Validation work shall be done by a team of experts drawn from various departments such as engineering, production, quality control and quality assurance. - Time constraints: Validation work shall be subjected to rigorous time schedules. These studies shall always be the last stage prior to taking new processes, facilities into routine operation. - Costs: Validation studies may incur cost as they require time of highly specialized personnel and expensive inputs. |  |  |  |
| 2.4 | Purpose:- |  |  |  |
|  | Validation Master Plan is a document, which describes our company’s intensions and the methods which are related with validation of the equipment, systems, utilities, facilities, material, analytical methods and process. Validation program is designed to demonstrate that the facility for the production up to final stage of production of different dosage forms is capable of meeting the process parameters in a repeatable and controllable manner. Validation master plan ensure that validation activities are carried out as per respective protocols and after completion will determine whether the equipment, system, process and method. - Meets the specifications of its designed. - Suitable for its intended applications. - Confirm to the basic cGMP design criteria. - Will satisfy the regulatory requirements. - Meets safety requirements as applicable. - Is capable of consistently producing a product that is fit for use. The critical utilities, Equipment and process validation program are established in accordance with the methods and procedure maintained by the product requirements which are based on the currently available product information and the Current Good Manufacturing Practices, guidelines and other regulations. The validation master plan helps; - The management to know and access what the validation programmed involves and understands its necessity. - The validation team members to understand their tasks and responsibilities. - To understand the company’s approach towards validation and the setup of organizing validation related activities - To defines the rational for performing validations versus the implementation of verification activity. - The plan also defines the process requirements in terms of test criteria, calibration and maintenance requirements, cleanliness requirements etc. |  |  |  |
| 2.5 |  | Scope:- - The scope of this document is to describe the systems and methodology used to execute the various phases of the validation program. - it applied to all critical equipment, instruments, procedures, utilities, facilities and other quality supporting systems used for manufacturing processing testing, labeling and packaging which may affect the quality of the product directly or indirectly. - Validation master plan applies methodology of validation program of the following. - Facilities. - Utilities (e.g. HVAC, water system and Compressed Air). - Control systems (e.g. computer hardware and software). - Manufacturing processes - Equipment/instruments/system(Qualification) - Cleaning processes - Analytical methods - Environmental (physical and Microbiological) - Personnel(Analyst, Checkers on Inspection or Packing line) - Revalidation or Re-qualification. All validation activities relating to critical technical operations, relevant to product and process controls within the firm shall find inclusion in VMP. This includes qualification of critical manufacturing and control equipment. |  |  |
| 2.6 |  | Validation team members:- Validation Team Constitution The overall validation process is co-ordinate by a team, which is a working group comprising of qualified personnel as required for the specific validation activities. |  |  |
|  |  | The validation team is made up of representatives from the following areas: - Quality assurance department - Production department - Engineering department - Quality control laboratory |  |  |
| 2.7 |  | Validation team is responsible for: - Preparation of validation Master plan - Determining the equipment’s, instruments, systems, facilities and utilities to be validated. - Preparation of validation and qualification protocols. - Execution of the validation and qualification protocols. - Determining the suitability of the established protocol to qualify a system. - Approving all validation protocols and report for each validation from respective functional department heads. - Reviewing each protocols and reports content to assure compliance with current regulations. - Verifying the adequacy of the tests used for the verification of the system as established by the validation protocols. - Determine the extent of revalidation necessity in case of changes in a validation system. - Document control. Other personnel will be utilized, as necessary, by member of the validation team to assist the assessment of the special equipment, system, utilities where their area of expertise will facilitate the development and execution and assure the quality of the validation efforts |  |  |
|  |  | Validation is a team effort that generally requires the involvement and close interaction of all therefore mentioned departments. Qualification Team Process-Validation Team Cleaning Validation -Team External Validation Services |  |  |
| 2.7.1 |  | Training requirements:- |  |  |
|  |  | Each member of the validation team shall be trained by qualified personnel on the validation process, protocol and SOP for all the validation activities. |  |  |
| 2.8 |  | Key elements of validation planning:- |  |  |
|  |  | Acceptance Criteria Fulfilled? NO Approval of reports Review -Team NO Review -Team Check of Test Specification YES NO Edit Test Specification? Repeat Test Procedure? Validation -Team Carry out the Programme OK? Testing & Programme Review -Team Editing Validation Team Validation Team Work Out the Programme YES |  |  |
| VALIDATION MASTER PLAN |  |  |  | DATE: May, 2019 |
| 2.9 |  | Testing and approving procedures |  |  |
|  |  | The team for VMP shall consist of personnel in the following positions: - Validation team (Review team): QA Manager Validation representatives - Validation representative: Person appointed by the validation team - Qualification/validation team: Personnel from different manufacturing area. (personnel from external validation services may be included, if required) - External validation services: External qualified personnel who support the qualification/validation team if necessary. |  |  |
| 3.0 |  | Responsibilities of each department: |  |  |
|  |  | Quality Assurance Department: A validation team member from QA department shall coordinate the entire validation process by discussing with the validation team, preparing validation protocols, monitoring the validation process, compiling and analysing validation data and test Results, and preparing the final report. Co-coordinating all change control approvals. Organizing the changes required to approve protocols resulting from specification or changes in operating parametersAll documentation associated with validation shall be reviewed by Head of Quality Assurance for completeness and compliance with cGMP requirements. Production Department: A validation team member from the production department shall be responsible to perform validation steps during manufacturing processes, equipment qualification and packaging processes. Providing all current and approved standard operation procedure to be used as part of manufacturing / validation process. Training of all manufacturing personnel in technical, validation and GMP aspects. Preparing reports on any deviation from occurs during process. Operating and maintaining plant, equipment, support systems and the specific manufacturing processes within its design limits.Production department shall prepare the packaging process and assist in the collection of validation data. Engineering Department: A validation team member from this department shall participate in performing validation; define the necessary equipment specifications, limitations, capacity, calibration and maintenance requirements; and provide necessary training on the proper operation and maintenance of the equipment. The engineering department shall be responsible for providing the necessary utilities and equipment accessories. Sometimes engineering department shall be required to participate in equipment installation and operational qualification. Quality Control Department: A validation team member from quality control department shall be responsible for Providing the necessary support for sampling, testing and reporting of test results for Validation. Participating in preparation of draft validation protocols. Training of all Quality control personnel in technical, validation and GMP/GLP aspects. Testing of Validation sample as per validation protocols and specifications. Preparing laboratory reports by collecting data and verified summaries of findings and conclusions. Ensuring that the laboratory systems, instruments, devices and equipment to be used in the course of validation studies have been qualified validated and calibrated. A support group in QC shall perform the microbiological testing and environmental monitoring during the validation process. |  |  |
| VALIDATION MASTER PLAN |  |  |  | DATE: Sept, 2022 |
| 4.0 |  | Concept of Qualification/Validation |  |  |
| 4.1 |  | Fundamentals:- |  |  |
|  |  | The aim of Qualification/validation of production facilities and process in BIOMEDICAL Limited, Ilorin (Nigeria) is to establish and provide documentary evidence that: - The facilities, equipment, and processes have been designed in accordance with the requirements of cGMP. - The facilities and the equipment have been built and installed in compliance with their design specifications. - The facilities and equipment operate in accordance with their design specification repeatedly and reliably produce a finished product of the required quality. A complete program for validation shall be developed for BIOMEDICAL Limited, Ilorin (Nigeria) to fulfil above requirements. Quality management shall provide that the entire products manufactured in the validated plant are suitable for use. Fundamentals of validation program:- The validation program shall be divided into five phase as follows, - Phase I – Preparation and Approved of the validation master plan. - Phase II – Preparation and Approved of the validation protocol. - Phase III - Execution of validation protocols, test laid down in validation protocols and data collection. - Phase IV – Preparation of validation Report, Compilation of data collection during validation, review, recommendations and Approval of validation Report. Document shall be kept under custody of the Quality Assurance. - Phase V – Ongoing evaluation, review, Change Control, Deviations and Revalidation. - The validation Master plan (VMP) shall be prepared by the validation Executive. - The document shall be checked by the heads of all functional area. |  |  |
|  |  | This VMP covers following fields of activities: - Qualification of Buildings and Area - Qualification of equipment and service units - Qualification of supply systems - Validation of Production processes - Validation of Cleaning processes - Periodic revalidation within the scope of change control plan. |  |  |
| 4.2 |  | Concept of validation life cycle |  |  |
| Validation Master Plan Revalidation Monitoring Maintenance Change Control DQ IQ OQ PQ Risk Analysis Process Validation VALIDATION REPORT Cleaning Validation 4.2.1 |  | Validation life cycle DQ: Design Qualification IQ: Installation Qualification OQ: Operation Qualification PQ: Performance Qualification |  |  |
| 4.2.2 |  | Actual survey of facilities, process, and process controls Validation Master Plan DQ Risk Analysis IQ OQ PQ Process Validation Cleaning Validation VALIDATION REPORT Revalidation Monitoring Maintenance Change control DQ: Design Qualification IQ: Installation Qualification OQ: Operation Qualification PQ: Performance Qualification |  |  |
| 4.3 |  | VALIDATION TYPES VALIDATION A Facility Validation B Analytical Method validation C. Vendor Qualification D. Control System Validation E. Equipment Validation Qualification 1. Design Qualification 2. Installation Qualification 3. Operational Qualification 4. Performance Qualification F. Process Validation G. Cleaning Validation H. Utilities Validation 1. Prospective 2. Concurrent 3. Retrospective 4. Revalidation 1. HVAC/AHU system |  |  |
| 4.3.1 |  | Facility Validation - Room number and location - Purpose of area for use - Dimension of area: Height (Meter), Length (meter), width (meter), Area (m2/ft2) Volume (m2/ft2). A )Area Details: - Walls: wall to floor, wall to ceiling, wall to return air riser, wall to door frame. - Ceiling : Type, Finish, Painting shade and Joining - Door : Type, Dimension (W X H meter), number and type of door closer, number and type of handle, size of glass view panel, Door interlocking. - Windows: Type and number of windows, Dimensions (H X H meter), Jointing. - Drainage: Type of trap, Number and location, Material of construction. - Electrical work: Wiring type, number of light fittings and type arrangement for emergency light. Number of intercom Point, Number of computer points, Number of power Sockets, Number of single phase sockets and type, number of three phase sockets and type, Number of lighting switch box and type illumination level in lux. Building management system: Indication box, size, location indication, pressure indication, temperature indication, and humidity and Door status |  |  |
| 4.3.2 |  | Analytical Method Validation: This type of validation shall be done at Quality control laboratory. |  |  |
| 4.3.3 |  | Vendor Qualification: - Vendor Qualification shall be carried out for all vendors of raw and packaging materials, to ensure that the components of the product, consistently meet predetermined specifications. - Vendor evaluation shall be done on the basis of questionnaire, pre shipment samples, machine trials and/or a vendor audit. - The supplied materials shall be checked /analyzed with respect to in-house specification. |  |  |
| 4.3.4 |  | Control System Validation: Computer software and hardware validation: - The objective of the computer system validation is to demonstrate that the computer systems (modify, archive and retrieve) as it is supposed to. - This shall be based on the following; - Specification for hardware and software depending on intended use. - Computer and manual modes of data validation giving similar results. - Levels of authority for data entry and change. - Accuracy and reliability of data entered. - Security system to ensure any alteration or pilferage of data. Detail description of control system - Control panel - Power supplies - Temperature sensor (wherever applicable) - Humidity sensor(wherever applicable) - Smoke detectors (wherever applicable) |  |  |
| 4.3.5 |  | Equipment validation (Qualification) - Design Qualification - Installation Qualification - Operation Qualification - Performance Qualification |  |  |
| 5.0 Elements of Qualification / Validation |  |  |  |  |
| 5.1 |  | - Design Qualification (DQ): |  |  |
|  |  | The purpose of the Design Qualification is to establish documentary evidence that the equipment design meet the process requirement for the manufacturing of Pharmaceuticals formulation at BIOMEDICAL Limited, Ilorin (Nigeria). The design qualification contains the design of equipment required for manufacturing of product and should agree with the design specification of customer. In the design qualification process, all qualified personnel from various departments of the company shall be involved. |  |  |
|  |  | Design Qualification (DQ) :- (Continue...) The Design Qualification is divided into the following sections: - Protocol:- - Design qualification report (report shall be prepared as per content provided in IQ-Protocol) - Approval procedure The person responsible from project department shall do the DQ and it shall be reviewed and checked by person responsible from production and maintenance department respectively. Finally, design Qualification docket shall be approved by QA Manager. 4)The following sections are to be included in Design qualification: - User requirement specification (URS): The following sections are to be included in user requirement specification: - Introduction - Overview - Operational Requirements - Capacity - Specification - Functional Requirements - Environment - Maintenance Life Cycle - Preferred vender list - Development - Testing - Delivery - Support Glossary References Approval - Functional Requirement Specification (FRS): - Specification provided by the supplier end is to be attached to the report. - Factory Acceptance Test procedure (FAT) - Remark of team of selector on fat - Approval |  |  |
| 5.2 |  | Installation Qualification (IQ):- The purpose of the installation qualification is to establish documentary evidence that all parts of equipment are installed according to the process requirements for the manufacturing in BIOMEDICAL Limited, Ilorin (Nigeria). It is documented that the operating criteria for the equipment, as installed are in compliance with the plant functional specifications. The IQ refers to as possible engineering documents (e.g., equipment drawing, calibration of equipment, plant functional specifications etc.)To avoid redundant documentation and to minimize overheads of updating. It may not be possible to inspection and verify every feature association with a piece of equipment. In such cases, a certificate of conformance shall be consider, obtained from the equipment manufacturer. The Installation Qualification is divided into the following sections: - Protocol:- - Installation Qualification report( report shall be prepared as per content described in the IQ-Protocol) - Approval - The following sections are to be included in protocol of installation Qualification protocol. - Approval - Objective - Scope - Responsibility - Procedure - Re – qualification criteria - Abbreviations - Reference |  |  |
| 5.3 |  | Operational Qualification (OQ):- The purpose of the operational qualification is to establish documentary evidence that the equipment and its accessories are capable of operating as per stated ranges and designed criteria. To start the OQ process for given equipment, it is necessary that the IQ process for same equipment has been completed. The witness shall be available by signature to ensure that the test procedure carried out by the operating person is in accordance with their test specifications. OQ demonstrates that the equipment/ Instrument/ System can operate throughout its dynamic range of operation, within limits and tolerances as per specified in the Operating manual or instructions. This demonstration of equipment’s basic operation shall involve all control Parameters of the equipment. OQ shall be carried out for equipment/ Instrument/ System that will be used routinely in manufacturing, testing or in critical utilities. The operational qualification is divided into the following sections: - Protocol - Operational Qualification report (report shall be prepared as per content described in the OQ-Protocol) - Approval The following sections are to be included in Operational Qualification protocol: - Approval - Objective - Scope - Responsibility - Procedure - Re-qualification criteria - Abbreviations - Reference |  |  |
| 5.4 |  | Performance Qualification(PQ):- The purpose of the performance qualification is to establish documentary evidence that the product manufacturing by using the equipment shall give consistent quality of product that meets the predetermined specifications. The PQ includes critical process parameters and critical quality attributes. To start the PQ process for given equipment, it is necessary the OQ process for same has been completed. The witness shall be available by signature to ensure that the operating and test procedure carried out by the operating person is in accordance with their teat and operating specifications. The performance qualification is divided into the following sections: - Protocol - Performance Qualification Report (report shall be contain certificate of analysis, yield of product, critical parameters etc.) - Approval - The following sections are to be included in performance qualification protocol: - Approval - Objective - Scope - Responsibility - Procedure - Acceptance criteria - Re-qualification criteria - Abbreviations - Reference |  |  |
| 6.0 |  | Process Validation (PV):- The Purpose Of Process Validation Is to provide documentary evidence that the process within their specified design parameters Is capable to produce consistent quality product of predetermined Quality. Also process validation is to demonstrate the reliability and reproducibility of manufacturing process, within established parameters and to assure batch uniformity and integrity of drug products. A process validation program shall establish a documented evidence of three full-scale production batches All Qualification Procedures For Equipment And Services Of The Manufacturing Process Must Be Completed Before Starting Process Validation Process. - The Process validation plan shall covers the following features: - Which manufacturing process must be validated (Product matrix) - Which kind of validation (Prospective, or concurrent) must be done - Schedule of the validation of product process - Number of batches that must be used for validation (normally 3batches shall be taken for Prospective and Concurrent validation) - Responsibilities - The Process validation program shall also contain the following key aspects: - Responsibilities - History of development and a description of the product (if available) - A manufacturing procedure and flowchart of the manufacturing process - A list of Raw material - A list of all equipment required for production. - Utilities - A list of production stages that may be critical for product quality - A list of quality parameters - A schedule for Process validation test procedures - Sampling procedures - Specification finished product - Acceptance criteria Process validation protocol shall be prepared before starting of Process validation, covering all above points. Process validation report shall be prepared as per content described in protocol. After completion of Checking and Reviewing procedure, Q.A Manager shall state his conclusion in the document. There are four types of the process validation. - Prospective validation - Concurrent validation - Retrospective validation - Revalidation Prospective Validation - The Prospective validation is used to establishing documented evidence prior to process implementation that a system does what is proposed to do based on pre-planned protocols. - Prospective Validation shall be carried out on the small scale or R & D batches during development stage. These validations were carried out to identify the processing steps, equipment settings and process timings, which can prove to be critical in overall impact of product quality. All these steps, equipment setting and process timing are evaluated to finalize the setting parameter. Limits of all critical parameters shall put tentatively. This draft form of study can lead towards the finalization of Batch Manufacturing Record / Batch Packing Record for large scale batches. The prospective validation shall take care of following: - Process description - Investigation summary of critical processing steps - Finished product release specification - Calibration of equipment - Appropriate analytical methods - Proposed in process control with acceptance criteria - Sampling plan Each prospective validation step shall be described in Qualification/Validation documents. Concurrent Validation:- - This validation comprises of determination and evaluation of process parameters applicable from scale up to large scale batches. - Process validation for establishing the predetermined limits of all critical parameters by conducting three consecutive commercial or large scale batches. - Validation carried out during routine production or products manufacturing for sale. - The tests are finalized based on evaluation of the results of these batches. These batches were monitored for stability and quality trends. - Matrix approach can be applied to process validation. The matrix approach means that a plan to conduct process validation on different strength of same product, manufactured by the same process and similar type of equipment. E.g. validation on one strength (preferably smallest) as a representative amongst the multiple strength of the same product. This approach should be in accordance with cGMP and should demonstrate that the process is consistent for all the strength. Note: The batches will not be dispatched unless and until the analysis of the validation sample is completed. Each concurrent validation step shall be described in Qualification/Validation documents. In these documents, the test methods for validation and acceptance criteria for the results are specified. Retrospective validation: - Historical trending of all critical parameters obtained from analytical results and / or IPQC results observed after testing of physical, chemical and microbiological parameters of product manufactured by same process. To ensure that all the parameters shall be in the limit and in uniform range. - Retrospective validation shall done yearly of products or batches manufactured in previous year. - Minimum first 20 batches shall consider for retrospective validation. If less than 20 batches were manufactured in one year then batches of previous years were consider for trending. Revalidation: The purpose of Revalidation is to maintain the validated status of plan, equipment and manufacturing process. - Periodic revalidation: This type of revalidation shall be done on periodic basis as per specified frequency. The revalidation shall be governed by the Change Control procedure. Revalidation after any change: All validated equipment, systems, processes, products etc. shall be continuously monitored for any change, to be considered for revalidation of the same. |  |  |
|  |  | Revalidation / Requalification [Verification of validation] shall be considered in the case of : ( A) Equipment / Instruments / System: - Substitution of existing equipment with a new equipment /instrument/System. - Change of site /location of the equipment / instrument / System from qualified existing place to other place. - Any major modification in the existing equipment / instrument /System since purchase which can have an adverse effect on the quality and efficacy of the finished product. ( B) MANUFACTURING PROCESS: - A change in the validated manufacturing process that has an effect on the quality and efficacy of the finished product, which includes a change in the following : - Master formula or manufacturing process. - Raw Materials - Equipment - Batch Size - Change in manufacturing site or location. - Change in the vendor of critical raw material - Change in the packaging material - Change in the cleaning processes or agents - Change which can affect the quality and effectiveness of product The scope of the revalidation procedures shall depend on the extent of the changes and the effect upon the product. |  |  |
| 7.0 |  | Cleaning Validation (CV):- The purpose of cleaning validation Is to provide documentary evidence that ensures that the cleaning procedure being used is suitable and delivers consistently, results for cleaning of equipment to obtain an acceptable level of cleanliness. Objective: - Cleaning validation is an important mechanism to protect the pharmaceutical products from cross contamination. A pharmaceutical product can be contaminated by previous product residues, traces of cleaning agent, micro-organisms or any foreign particle as borne particles, dust, raw materials etc. - This should be performed to provide a documented evidence that the procedure being followed for cleaning of equipment and accessories is effective and removes residues of previous batch / product up to a predetermined acceptance level, using a well-defined protocol and acceptance criteria. - For equipment which are not product specific, cleaning procedure for such equipment should be able to ensure that it can consistently remove residues to the predetermined level. Procedure: - There are two methods employed for sampling during cleaning validation. These are Rinse and Swab method. - Rinse method will be applicable when swabs are not practically possible. - Cleaning procedures for products and processes, which are very similar are do not need to be individually validated. - A single validation study under consideration of the “worst case “can be carried out on the basis of the maximum strength, potency, solubility and LD. The cleaning validation can be evaluated, - For highest strength product in multi strength products or all products. - For highest potency product in all products. - For least solubility in common cleaning solvent i.e. Water. - This “worst case” term can also be referred as “bracketing”. And such approach is called as matrix approach. - Three cleaning validations exercise shall be performed and shown to be successful in order to prove that the method is validated. - The documentation of cleaning validation shall include validation protocol, cleaning procedures, sampling plans, surface area calculations, testing procedures and executed record of the cleaning event, a record of the sampling and testing events and a validation report which contains clear conclusions against acceptance criteria. - This validation shall covers sanitation, microbiology as well as removal of external agents like detergents, using sampling techniques capable of detecting and measuring to very low levels wherever applicable. - There are two criteria’s for the cleaning validation, one is dose criteria and second is 10 PPM criteria. - Validated analytical methods used for cleaning validation shall have sensitivity to detect residues or contaminant of previous product. The detection limit for each analytical method should be sufficiently sensitive to detect the established acceptance level of residues or the contaminant. Equipment surface area calculation: - The product contact surface area for each equipment shall be calculated by theoretically breaking down the equipment in to basic geometric shapes [cylindrical, cubical, cone, rectangle etc.] and then apply to relevant area calculation formulae. - Amount of contaminant, mix from previous product to next product, depend on the common surface area of equipment used for both previous and next product. The Acceptance Criteria For The Cleaning Validation: - No residue shall be visible on the equipment surface after cleaning procedures are performed. - No more than the limit specified of the residue shall remain as residue in any equipment after cleaning has been done. Total residue of all equipment observed In rinse sample shall be less than the specified limit. - Total residue of all equipment observed in swab sample shall be less than the specified limit. - First batch of the next product taken after the cleaning procedures are performed shall pass predetermined specifications. - At least three consecutive cleaning runs / opportunities shall meet the above mentioned acceptance criteria before the cleaning procedures can be considered as validated. - For multi-purpose/dedicated equipment, it is acceptable to visually confirm that the equipment is clean after batch-to-batch cleaning procedures. Therefore, following consideration towards cleaning validation for such equipment may not be necessary. The cleaning validation plan shall contain, for example, which manufacturing processes must be validated and which kind of validation (prospective, or concurrent) must be done - Schedule Of The Validation Of Product Processes - Number Of Batches That Must Be Sued For Validation (Normally 3 Batches Shall Be Taken For Prospective And Concurrent Validation) - Responsibilities - The Following Points Shall Be Included In The Cleaning Validation Program. - Purpose Of Validation - Responsibilities - Contents Of Sops And Cleaning procedure - Method validation of test procedure - Methodology of Cleaning validation - Sampling Procedure - Acceptance Criteria |  |  |
| 8.0 |  | Utilities Validation: Equipment / Machineries used for plant utility as, - HVAC system - Water system - Pure steam generator - Compressed air system HVAC system - Air Handling Unit system, ventilation system, and evaporating system are the part of HVAC system which is one of the utility have directly or indirectly impact on products. - Total numbers of independent air handling units (10), ventilation systems with supply and exhaust duct have been provided in the facility. Each critical operational area has a dedicated air-handling unit (AHU), Magnehelic gauges and temperature and humidity controls. Each room is provided with terminally filtered air, as per the requirement of operation, to be carried out in the room. The core corridors are provided with controlled environment. Pressure balancing of each room has been done to avoid any cross contamination, by any ingress of air / drug product from adjacent areas. - There is provision to automatically control and record temperature and percentage relative humidity & monitor and record the pressure differential across Rooms - This system is qualified initially, monitored at regular intervals and external validation/calibration is done twice in a year. - Initially Temperature control test, Humidity control test, Air velocity & number of air changes test, Differential Pressure Control test, Air borne Particle Count test, Microbial count test, Air flow direction test shall be conduct to perform the system. - Limits for each test shall mention in the validation protocol as acceptance criteria. Acceptance criteria: - The system should be installed in accordance with design specifications, manufacturer recommendation, and cGMP - Instruments should be calibrated, identified and entered into the calibration program - General controls and alarms should be operated in accordance with design specification - The system should be operated in accordance with design specification throughout the operating range or range of intended use - System regulators must operate within predetermined operating pressure - The system capacity must be sufficient to operate all systems at peak demand periods |  |  |
| 9.0 |  | Water system validation - Water system is one of the utility which have directly or indirectly impact on products. There are mainly four types of water system as follows, - Raw water system [Potable water] - Purified water system - Reverse Osmosis water system - MCDP water system - New water system shall be established and use for regular purpose after one year validation study includes the results of chemical and microbiological quality of the water. For qualifying the water system, Chemical ,Microbiological Test and Pyrogenicity test shall be conduct mentioned as follows; - pH, Conductivity, Colour, Total hardness, Chloride, heavy metals, Microbial counts, pyrogens and endotoxins. Test method for each test shall be mentioned in the protocol. - Location of sampling points shall be done as per specified in Validation protocol or as per SOP. - Frequency for all above mentioned tests shall be mentioned in the individual validation protocol. - Validation of water system shall divide in to three different phases as follows. Frequency of sampling, number of samples taken and no. of sampling locations are reduces from phase I level to next phase level. The Performance Qualification of water system shall be carried out in three phases: - Phase-1 – extensive sampling from Generation points & usage points for 4 weeks - Phase-2 – extensive sampling from selective generation & all usage points for 3 months - Phase-3 – routine sampling from usage points on rotation basis, for 8 months Following acceptance criteria: - Raw water circulation and its storage shall be requalifies if microbial load is more than the specifications. Chemical water quality shall conforms with WHO/In-house standards. - Purified water circulation and its storage shall be qualified if microbial load is if, not more than 100 CFU / ml and chemical water quality shall conforms with WHO /In-house standards. - Reverse Osmosis water system and its storage shall be qualified if microbial load is if, more than 10 CFU / ml and chemical water quality shall conforms with WHO /In-house standards. - MCDP water system and its storage shall be qualified if microorganism is NOT PRESENT in the water and chemical water quality shall conforms with WHO /In-house standards. - Sanitization of water system shall be done after 15 days - Water system shall be revalidating if there is any major change, replacement or modification in existing system. It shall be periodically revalidate or verified once in a year on the basis of monthly, weekly or daily testing data as per scheduled trend. ACCEPTANCE CRITERIA The specific acceptance criteria for the experimentation results, during validation shall be defined in detail in the respective validation protocol. Recognized national or international standards shall be cited for reference. |  |  |
| 10.0 |  | Validation Report The validation report refers following sub-sections: - Qualification : Includes DQ, IQ, OQ, PQ - Process validation : PV - Cleaning validation : CLV - Water Validation : WV - Analytical method Validation : AMV Contents: A detailed discussion of the following points shall be included in validation report - Changes and deviations of the qualification /validation activities from the documentation shall be listed. The measures taken instead must be justified and incomplete measures must be listed. - Critical points found during qualification/validation shall be listed and the measures taken described. - The validation report shall include overview of all documentation including annexure. At the end of the validation report, a conclusion shall be given as to whether or not the equipment and the process have the status “qualified /validated”. |  |  |
| 11.0 |  | Change of Documents If it is necessary to edit or remove test specifications from the documentation, this shall be done according to the following procedure. The version number of the test specification shall be increased by one after specification is edited. - The results of the test shall be updated. - The history shall be updated - The table content shall be updated The changes shall be approved by qualified personnel by signature on a new coversheet. |  |  |
| 12.0 |  | BUIDING DESCRIPTION |  |  |
|  |  | The design and material used for the construction of building is according to cGMP regulations. All walls, ceilings, and flooring inside the production area are designed as per cGMP regulations. The building is divided into two floors, - Production area (ground floor) - Service floor for utilities. |  |  |
| 13.0 |  | EQUIPMENT DESCRIPTION ( check the list of equipment Annexure) |  |  |
| 14.0 |  | Validation responsibilities: |  |  |
|  |  | The validation team shall be responsible for validating the facility. The QA Manager shall be responsible for approval of all validation protocols, final reports, SOPs, chemical and microbiological testing, as well as other critical cGMP documentation. |  |  |
| 15.0 |  | Design and review: |  |  |
| 15.1 |  | The validation of company’s facility shall start with the initial conceptual design. Design review include; - Architectural layout - Civil work execution - HVAC system - Water system - Personnel and material flow - Equipment - Material receipt, storage and dispatch - Quality assurance/control testing and inspection - Maintenance and calibration |  |  |
| 15.2 |  | Installation Qualification protocols |  |  |
|  |  | The installation qualification (IQ) verifies that the equipment or system and/or corresponding utilities are installed in accordance with design specifications, manufacturer’s recommendations, and cGMP. The IQ shall confirm that critical instruments are calibrated and that system components are properly identified. Any exceptions shall be documented, corrected and/or justified. Typical installation qualification test functions include, but are not limited to the following: - Manufacturer specification - Purchase order specification - Piping and installation drawing - Construction and installation - Test equipment calibration - Required spare parts - Cleaning /passivation - System installation compliance to CGMP The IQ shall be performed in accordance with a reapproved written protocol. |  |  |
| 15.3 |  | Operational Qualification Protocols |  |  |
|  |  | The operational qualification (OQ) verifies that the system or equipment operates in accordance with design specification, manufacturer recommendations, and cGMPs. OQ testing is designed to stimulate the full range of operating conditions to establish a system capabilities baseline and to ensure that the system operates as intended. |  |  |
|  |  | Typical Operational Qualification test functions include, but are not limited to the following: - General operational and control verification - Functional testing of computer-related systems - Specific operational testing - Utility system - Calibration of test equipment - Standard operating procedures(SOP) - Cleaning procedures etc. The Operational Qualification shall be performed in accordance with a reapproved written protocol. |  |  |
| 15.4 |  | Performance Qualification Protocols |  |  |
|  |  | The performance qualification (PQ) verifies that the performance of critical utility systems or processes is in accordance with the process requirements. Critical utility system such as purified water, air, shall be challenged throughout proposed operating ranges for extended periods of time, while an extensive program of quality monitoring shall be performed. At least three consecutive batches shall be taken for the performance qualification. Performance qualification shall be carried out in accordance with a reapproved written protocol. The specific PQ protocol shall be developed from the finished product specifications, cGMP requirements, and other specific documents. |  |  |
| 16.0 |  | Process Validation Protocols |  |  |
|  |  | Process Validation verifies the performance of the overall product manufacturing process. Process validation shall be performed on the entire product manufacturing process, which shall include all support, processes, packaging and labelling, critical process parameters, critical quality parameters, utilities etc. During process validation, three consecutive batches shall be manufactured and all in-process and finished product specification shall be verified. Critical parameters shall be set at different ranges during the manufacturing process to verify that the product consistently meets its predetermined quality specification over the range of critical parameter settings. All equipment processes, and testing procedures must be validated prior to the start of process validation. Process validation shall be performed in accordance with a pre-approved written protocol. The specific PV protocol shall be developed from finished product specification, the manufacturing process capability, cGMP requirements and other specific documentation. These along with the acceptance criteria shall be approved by Head QA. |  |  |
| 17.0 |  | Validation master report |  |  |
|  |  | The final report shall summarize the result of the validation process and provides an analysis of the test data in support of the conclusion that the equipment or system demonstrates consistent performance within the established acceptance criteria. Deviations and/or exceptions to approved protocols, along with suitable explanations and justifications, shall be documented in the final report. |  |  |
| 18.0 |  | Validation package |  |  |
|  |  | The final reports, completed protocols, and all supporting documentation shall be placed into final report package for checking, reviewing and approval. Deviations and/or exceptions to approved protocols, along with suitable explanations and justifications, shall be documented in the final report. |  |  |
| 19.0 |  | Certification for use in manufacturing |  |  |
|  |  | Validation of a system /equipment/process shall be considered complete when its respective protocols identified in the validation master plan have been successfully implemented and the final report has been approved. A system /equipment/process shall be certified for use in manufacturing upon completion of the following: - Applicable protocols have been successfully completed and the final report has been approved. - Applicable SOPs have been approved. - Validation shall be completed for all utility systems that support the equipment or process, and for all equipment that supports a process - All system /equipment/process shall be under calibration criteria |  |  |
| 20.0 |  | Required protocol and procedures for production |  |  |
|  |  | S.No; System /Equipment; Protocols; SOP’s. DQ; IQ; OQ; PQ; CL; OP; PM; CL; CAL. 01; Demineralization Plant; √; √; √; √; √; √; √; √; √. 02; Cleaver Boiler 5Tons; √; √; √; √; √; √; √. 03; Atlas Copco Air Compressor; √; √; √; √; √; √; √. 04; Power flex Air Compressor; √; √; √; √; √; √; √. 05; 132 KVA Generator; √; √; √; √; √; √; √. 06; 450 KVA Generator; √; √; √; √; √; √; √. 07; 800 KVA Generator; √; √; √; √; √; √; √. 08; Reverse Osmosis machine; √; √; √; √; √; √; √; √; √. 09; Multi-Column Distillation Plant; √; √; √; √; √; √; √; √; √. 10; Jacketed WFI Tank(6000 Litres); √; √. 11; 3 Stainless steel Mixing tanks (2000 litres); √; √; √; √. 12; Digital Weighing Scale 15Kg; √; √; √; √; √; √; √. 13; Digital Weighing Scale 50Kg; √; √; √; √; √; √; √. 14; Batch Coding Machine; √; √; √; √; √; √; √. 15; Semi Automatic Filling Machine; √; √; √; √; √; √; √; √. 16; Automatic Filling Machine; √; √; √; √; √; √; √; √. 17; De lama Autoclave; √; √; √; √; √; √. 18; Phamalab Autoclave; √; √; √; √; √; √. 19; Analytical weighing Balance; √; √; √; √; √; √; √; √; √. 20; UV-VIS Spectrophotometer; √; √; √; √; √; √; √; √; √. 21; Polarimeter; √; √; √; √; √; √; √; √; √. 22; Flame Photometer; √; √; √; √; √; √; √; √; √. 23; PH Meter; √; √; √; √; √; √; √; √; √. 24; Incubator; √; √; √; √; √; √; √; √; √. 25; Conductivity meter; √; √; √; √; √; √; √; √; √. Here, DQ: Design Qualification IQ: Installation Qualification. OQ: Operational Qualification PQ: Performance Qualification CL: Cleaning Procedure CAL: Calibration Procedure FR: Final Report OP: Operating Procedure CV: Computer Validation PM: Preventive Maintenance Procedure |
| 21.0 |  | Calibration program summary |  |  |
|  |  | Calibration is defined in related SOPs - All process and facility instruments are classified to their calibration status into one of the three following categories: - Calibration - Preventive Maintenance - No calibration/ no preventive maintenance - All instruments providing critical process information necessary to make a quality documentation shall be calibrated with the available standards. - User department is responsible for scheduling, tracking and maintenance of standards, record, etc. - All instruments providing critical process information shall be calibrated before performing the validation studies. List of equipment accordingly: - : Calibration; Preventive maintenance. Pressure gauge Vacuum gauge Magnehelic gauge Conductivity meter Temp controller Temp. Indicator Temp. Ind. Controller with sensor Thermometers Mercury glass thermometer Timer Mixing timing PH meter Conductivity meter Weighing Balance Lab. Equipment; Weighing scales RO. Machines DM plant Air compressors Batch coding machine Autoclaves Transfer pump HVAC Mixing tank Filling machine. |
| 22.0 |  | PREVENTIVE MAINTENACE PROGRAM SUMMARY |  |  |
|  |  | An appropriate preventive maintenance program shall be defined for the equipment that shall be checked at the time of validation The in-charge of the related area and the engineer concerned shall sign the register and verify if repairs are required. |  |  |
| 23.0 |  | KEY STANDARD OPERATING PROCEDURE (SOPs) |  |  |
|  |  | List of key procedures for process operations and maintenance are as follows: - Raw material issuing system - Packaging material issuing system - Manufacturing procedures - Batch packing record system - Equipment operating procedures - Equipment calibration procedure - Equipment maintenance procedure - Product release procedure - Handling out of specification result - Complaints handling system - Change control system - Quality Risk Assessment |  |  |
| 24.0 |  | VALIDATION OF BUILDING |  |  |
| 24.1 |  | Civil work |  |  |
|  |  | Test functions: - Inspect the facility as per approved layout - Inspect floors in each room - Inspect painted surface in each room - Perform dimensional testing for each room - Inspect lights in each room - Inspect utilities into each room - Inspect all drains - Inspect all doors and interlocks |  |  |
|  |  | Acceptance criteria: - The facility is constructed in accordance with design specifications and cGMP. - Doors are opened/closed properly, and interlocks operates in accordance with design specification - Floor surfaces are smooth, and free of holes and soft spots - Painted surfaces have full and smooth coverage - The length, width, and height of each room conforms to design specification - All light switches and light should be adequate at all work stations - All utilities penetration are to be correctly installed, labelled and sealed - All drains are located according to the design specification |  |  |
| 24.2 |  | Drainage system |  |  |
|  |  | Test functions: - Inspect the drainage system as per the approved layout - Inspect floors drain locations as per the approved layout - Check that process drains are segregated from the sanitary waste - Ensure that process drains are provided with deep-seal traps Acceptance criteria: - The system is constructed and installed in accordance with design specification, manufacturer recommendations and cGMP. - Open floor drains are not to be located at core areas - Process drains are to be separated from sanitary waste drainage - Process drains have adequate back-flow preventive devices |  |  |
| 25.0 |  | VALIDATION OF UTILITY SYSTEMS |  |  |
|  |  | This part of validation master plan provides a summary of the key validation test functions and acceptance criteria for each utility system. Water system (DM Water Plant) Test functions: - Test the water produced by the system to ensure adequate conductivity - Check that the holding tank water quality does not change adversely during storage holdings in tanks. |  |  |
|  |  | Acceptance criteria: - The system should be installed in accordance with design specifications, manufacturer recommendation, and cGMP - Instruments should be calibrated, identified and entered into the calibration program - General controls and alarms should be operated in accordance with design specification - The system should be operated in accordance with design specification throughout the operating range or range of intended use - System regulators must operate within predetermined operating pressure - The system capacity must be sufficient to operate all systems at peak demand periods. |  |  |
| 25.1 |  | Compressed air |  |  |
|  |  | Test functions: - Perform installation qualification - Perform general operational controls verification testing - Operate system throughout the range of operating design specification or range of intended use - Verify that the compressed air system is capable of supplying pressurized compressed air to all use points, perform an operational test of the distribution system and pressure regulators by monitoring the pressure output at the respective use points - Perform a capacity test to verify that the system is capable of supplying the required pressure, and flow rate at each use point - Verify that in-line filters are integrity tested. Confirm that all documentation clearly indicates acceptable test result - Perform viable particulate count and microbiological testing Acceptance criteria: - The system should be installed in accordance with design specification, manufacturer recommendation and cGMP. Instruments should be calibrated, identified, and entered into the calibration program - General controls should be operated in accordance with specification design - The system should be operated in accordance with design specification throughout the operating range or range of intended use - The compressed air distribution system must consistently deliver pressurized compressed air to the use point at the design pressure as specified - All in-line filters should be integrity tested and qualify per manufacturer’s operating specifications. |  |  |
| 26.0 |  | Purified water |  |  |
|  |  | Test functions: |  |  |
|  |  | - Perform installation qualification - Perform general operational controls verification testing - Verify that the purified water system is capable of supplying water to all use points, perform an operational test of the distribution system and pressure regulators by monitoring the pressure output at the respective use points - Perform a capacity test to verify that the system is capable of supplying the required water, pressure, and flow rate at each use point - Verify that filters are integrity tested. Confirm that all documentation clearly indicates acceptable test result - Perform water analysis chemical as well as microbiological - Check bio-burden if necessary Acceptance criteria: - The system should be installed in accordance with design specification, manufacturer recommendation and cGMP. Instruments should be calibrated, identified, and entered into the calibration program. - General controls should be operated in accordance with specification design - The system should be operated in accordance with design specification throughout the operating range or range of intended use - The purified water distribution system must consistently deliver water to the use point as per the design specified - All filters should be integrity tested and qualify per manufacturer’s operating specifications. - The microbiological test result must meet the approved specification requirements |  |  |
| 27.0 |  | QUALITY CONTROL LABORATORY VALIDATION The quality control laboratory serves one of the most critical functions in BIOMEDICAL Ltd. Consequently, a comprehensive validation program shall be initially performed for procedures and equipment used for testing of all products, as well as any analytical procedures. Thereafter, the analytical methods requirements and resulting validation shall be evaluated and performed as required for each new product. |  |  |
| 27.1 |  | Laboratory equipment qualification Laboratory equipments such as incubators, refrigerators, polar meter, etc. Shall be qualified using the same approach as that used for process equipment and utility systems. Validation of Facility/Equipment/Systems to be carried out Validation of Facility/Equipment/Systems includes designs, installation, and operational qualification. INSTALLATION QUALIFICATION: Test functions: - Checks for the fulfilment of the manufacturer’s recommendations (utilities such as electricity, water and environmental conditions such as humidity, temperature, vibration level and dust) and allow sufficient shelf space for the equipment, SOPs, operating manuals and software. - Describe vendor selection, instrument design parameter and its capacity in validation docket - Compare equipment, as received, with its purchase order (including software accessories, spare parts) - Check documentation for completeness (operating manuals, maintenance instructions, standard operating procedures for testing, safety validation certificates) - Checks equipment for any damages - Switch on the instruments and ensures that all modules power up and perform an electronic self-test. - Install software on computer following the manufacturer’s/supplier’s recommendation - Verify correct software installation, e.g., are all files loaded. Utilities to do this should be included in the software itself - Make back up copy of the software and configure peripherals, e.g., printers and equipment modules - Identify and make a list with a description of all hardware, include drawings where appropriate - List equipment manuals and SOPs. - Prepare an installation report. Acceptance criteria: - The system should be installed in accordance with design specifications, manufacturer’s recommendations, and cGMPs. Instruments should be calibrated, identified, and entered into the calibration program - The system should be operated in accordance with design specification throughout the operating range or range of intended use - The safety device should operate as specified in the manual OPERATIONAL QUALIFICATION: Test functions: - Calibrate the parameters like temperature, rpm, gauges, dimension etc. - Check functional parameters of operation - If calibration and functional operation is not satisfactory, get it rectified from the engineer and recheck the instrument - If calibration and functional operation is satisfactory, proceed for performance qualification. Acceptance criteria: - The system should operate as per testing requirement - Calibration and function operation must give satisfactory result. PERFORMANCE QUALIFICATION: Test function: - Carryout the performance of instrument as per standard calibration procedure of the respective instrument - Measure performance characteristic of instrument and compare with the document preset limit - If performance is satisfactory, instrument may be released for use - If performance is not satisfactory, call service engineer/vendor for necessary corrective actions - Record all the result and get approved - All the documents like purchase order, supplier’s validation certificates, list of manuals etc., to be attached with the report of performance qualification Acceptance criteria: - The system should perform in accordance with manufacturer recommendation - Instrument should be calibrated, identified and entered into calibration program. CALIBRATION: - Each equipment of quality control should undergo calibration program - Each equipment should be calibrated before use - Analytical weighing balance should be calibrated with standard weights daily - Prepare the calibration format for each instrument. All records should be stored in calibration file - If calibration results are not acceptable, rectify the error if any - Calibration frequency of QC ( SEE THE ANNEXURE) |  |  |
| 27.2 |  | METHOD VALIDATION: |  |  |
|  |  | - Each analytical method should be validated before use - Method validation should contain following performance parameters. - Precision - Accuracy - Selectivity & specificity - Limit of detection - Limits of quantification - Linearity and range - Ruggedness or robustness - Stability of analytical solution - Check following parameters during quantitative analysis of active ingredients and finished product - Accuracy - Precision - Specificity - Linearity - Range - Check following parameters during quantitative analysis of impurities in RM & finished products - Accuracy - Precision - Specificity - Check following parameters during quantitative analysis of finished products before cleaning validation. - Blank interference - Stability of analytical solution - Linearity - Range - Limit of quantification - Limit of detection CLEANING PROCEDURE |  |  |
|  |  | - Cleaning procedure used to clean the equipment should be validated - Analytical method used for testing of cleaning validation samples should be validated - Perform the test and sampling as described in cleaning validation protocol and specific SOP for the cleaning of such equipment - Equipment should be cleaned up to acceptance level of cleaning - Perform both rinse and swab sampling method - Prepare the cleaning validation report. |  |  |
| 14.2 |  | Computer related systems used in quality control laboratory - Computer related systems used in the laboratory should be validated - Data entry authorisation (e.g. password protected) - Security of the data base should be maintained. |  |  |
| 28.0 |  | Validation criteria and qualification frequency ( SEE THE ANNEXURE) |  |  |
| 29.0 |  | ANNEXURES ANNEXURE 1 - Product lists Annexure II – Equipment lists Annexure III - water process flow ( Infusion) Annexure IV - water process flow ( Oral Liquid) Annexure V – production process Flow ( Infusion) Annexure VI – Production process flow ( Oral Liquid) Annexure VII – Equipment Validation Schedule Annexure VIII – Process Validation Schedule |  |  |
| 30.0 |  | GLOSSARY 1. Acceptance criteria: The product, equipment, and /or process specifications and limits, such as acceptable quality level and unacceptable quality level, that is necessary for making a decision to accept or reject. 2. Calibration: Demonstrating that a measuring device produces results within specified limits of those produced by a reference standard device over an appropriate range of measurements. This process results in corrections that may be applied if maximum accuracy is required. 3. Change control: A formal system by which qualified representatives of appropriate discipline reviewed proposed or actual changes that might affect a validated status. The intent is to determine the need for action that would ensure and document that the system is maintained in the validated status. 4. Certification: Documented testimony by qualified authorities that a system qualification, calibration, validation or revalidation has been performed appropriately and that th9. Installation Qualification: Documented verification that all key aspects of the installation adhere to appropriate codes and approved design intentions and that manufacturer’s recommendations are suitably considered. 5. Operational Qualification: Documented verification that the system or subsystem performs as intended throughout all anticipated operating ranges. 6. Performance qualification: Documented verification that the equipment performs as intended for the process, as specified by operating ranges. 7. Process optimization: To optimize the process for maximum efficiency while maintaining quality standards. 8. Process validation: Establishing documented evidence that a process does what it purports to do. It is a documented verification that the integrated system functions as intended in its normal operating environment. The term performance qualification may also be used. 9. Pre-determined acceptance criteria: The criteria assigned before undertaking testing to allow evaluation of test results to demonstrate compliance with test phase of delivery requirement. 10. Protocol: A document that gives a direction or guideline for any process as what, where, when, who, why and how to do. 11. Qualification: The performance of tests to determine if a component of a manufacturing process possesses the attributes required to obtain a specified quality of a product. 12. Quality assurance: The activity of providing, to all concerned, the evidence needed to establish confidence that the quality function is being performed adequately. 13. Quality control: The regulatory process, through which industry measures actual quality performance, compares it with standards and acts on the difference. 14. Quality function: The entire collection of activities from which industry achieves fitness for use, no matter where these activities are performed. 15. Revalidation: Repetition for verification of the validation and qualification process or a specific portion of it. 16. Validation change control: A formal monitoring system by which qualified representatives of appropriate disciplines review proposed or actual changes that might affect validated status and cause corrective action to be taken that will ensure that the system retains its validated state of control. 17. Validation report: A scientific report of the results derived from executing a validation protocol. A brief summary of conclusions from a specific task conclusions report, usually indicating validation success and designating proven acceptable ranges that have resulted. The conclusions are formally approved. 18. Validation Master Plan: A document providing information on the company’s validation activities. It should be define details of and time scales for the validation works to be performed. Responsibilities relating to the plan should be stated. |  |  |