LOCATION – WTP/Lab

PROTOCOL No. – BML/QA/pHQ/001

Effective date: August, 2025.

PROVIDED BY:

BIOMEDICAL LIMTED

KWARA STATE, NIGERIA

BIOMEDICAL LIMITED

pH METER QUALIFICATION PROTOCOL

Installation Qualification / Operational Qualification / Performance Qualification

Contents

1.  Purpose	5

2.  Scope	5

3.  Regulatory and Reference Basis	5

4.  Instrument Description and Identification	5

5.  Roles and Responsibilities	6

6.  Pre-Qualification Requirements	6

7.  Installation Qualification (IQ)	8

7.1  IQ Acceptance Criteria	8

7.2  IQ Test 1 – Document Verification	8

7.3  IQ Test 2 – Physical and Hardware Verification	8

7.4  IQ Test 3 – Utilities Verification	9

7.5  IQ Summary and Sign-off	9

8.  Operational Qualification (OQ)	10

8.1  Buffer Standards Used	10

8.2  OQ Test 1 – Instrument Power-On and Self-Diagnostic	10

8.3  OQ Test 2 – Single-Point Calibration Verification (Neutrality)	10

8.4  OQ Test 3 – Two-Point Calibration and Linearity (pH 4.00 and 10.01)	11

8.5  OQ Test 4 – Electrode Slope Check	11

8.6  OQ Test 5 – Temperature Compensation Verification	11

8.7  OQ Test 6 – Response Time	11

8.8  OQ Test 7 – Reproducibility (Repeatability)	12

8.9  OQ Summary and Sign-off	12

9.  Performance Qualification (PQ)	13

9.1  PQ Test 1 – Inter-Day Precision (3 Days)	13

9.2  PQ Test 2 – pH Measurement of LVP Products	13

9.3  PQ Test 3 – Accuracy Against External Reference	13

9.4  PQ Test 4 – Operator-to-Operator Variability	14

9.5  PQ Summary and Sign-off	14

10.  Deviation Management	15

11.  Requalification Criteria	15

12.  Document Control	15

13.  Appendix	15

Appendix A – Buffer pH vs Temperature Reference Table	15

Appendix B – Glossary	16

PROTOCOL FORMAT APPROVAL:

Signatures on this page indicate that this document and its format have been prepared, checked and approved by the personnel concerned. It shall be signed and date as shown below

|  | PREPARED BY: | REVIEWED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

|  | APPROVED BY: | AUTHORIZED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

# 1.  Purpose

This protocol defines the Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ) requirements for the pH meter designated for use in the LVP Plant and Quality Control laboratory at Biomedical Ltd (BML). The document establishes documented evidence that the instrument is correctly installed, operates within defined parameters, and consistently performs accurate pH measurements suitable for GMP-regulated pharmaceutical manufacturing.

Successful completion of all three qualification phases, as documented in the companion Qualification Report (BML/EQ-RPT/PHM/001), is required before the instrument is approved for routine use.

# 2.  Scope

This protocol applies to the pH meter installed for in-process pH measurement, final product release testing, and cleaning validation support within BML's LVP Plant. It covers:

- The main pH meter unit and all integrated accessories (electrode(s), temperature probe, ATC module, calibration solutions)

- Any spare or replacement electrode of the same specification

- pH measurements performed on LVP products: Normal Saline 0.9%, Dextrose 5%, Ringer's Lactate, and CAPD solutions

This protocol does not cover the Oral Liquid Plant pH meter, which is qualified under a separate document.

# 3.  Regulatory and Reference Basis

| Reference | Applicability |
| --- | --- |
| EU GMP Annex 15 (2015) | Qualification and Validation framework |
| EU GMP Chapter 4 | Documentation requirements |
| ICH Q2(R1) | Analytical method validation – applies to instrument performance |
| USP <791> pH | pH measurement requirements for pharmaceutical products |
| Ph.Eur. 2.2.3 | Potentiometric determination of pH |
| WHO TRS 937 Annex 4 | WHO GMP for pharmaceutical products |
| NAFDAC GMP Guidelines | National regulatory requirements (Nigeria) |
| BML/SOP/QC/005 | Operation and Calibration of pH Meter |
| BML/CAL/REG/001 | BML Calibration Register and Schedule |

# 4.  Instrument Description and Identification

| Parameter | Details |
| --- | --- |
| Equipment Name | pH Meter – Laboratory/Process Grade |
| Manufacturer | [To be completed on receipt of equipment] |
| Model | [To be completed] |
| Serial Number | [To be completed] |
| Electrode Type | [e.g., Combined glass-reference, ATC-capable] |
| Measurement Range | pH 0.00 – 14.00 |
| Resolution | 0.01 pH units |
| Accuracy Specification | ±0.01 pH units (manufacturer's specification) |
| Temperature Compensation | Automatic (ATC), 0 – 100°C |
| Power Supply | [Mains / Battery – to be confirmed] |
| Location | LVP Plant – QC Laboratory / WTP |
| Calibration Interval | Daily (before use) / External: 3 months |

# 5.  Roles and Responsibilities

| Role | Responsibility |
| --- | --- |
| QA officer | Protocol authorship, test execution oversight, review of results, preparation of the Qualification Report |
| QA Manager | Review and approval of protocol and report; final release decision |
| Production Manager | Provision of installation location; operational approval |
| Plant Engineer / Maintenance | Physical installation verification, electrical safety checks |
| QC Analyst | Execution of OQ and PQ tests under the supervision of QAQC Head |

# 6.  Pre-Qualification Requirements

The following items shall be confirmed before commencing any qualification activity:

| Pre-Qualification Check | Pass Criterion | Status | Init. |
| --- | --- | --- | --- |
| Purchase Order and delivery documentation available | Documents on file |  |  |
| Instrument received undamaged – visual inspection complete | No damage |  |  |
| Manufacturer's Certificate of Conformance (CoC) / Factory Calibration Certificate received | CoC on file |  |  |
| Approved operating SOP (BML/SOP/QC/005) in place and accessible | SOP current |  |  |
| NIST-traceable reference buffer standards procured (pH 4.00, 7.00, 9.20 ±0.01 at 25°C) | Standards available |  |  |
| Environmental conditions at installation site within specification (Temp: 18–25°C; RH: 30–70%) | Within range |  |  |
| Dedicated analyst assigned and trained on SOP | Training record |  |  |
| Instrument registered in BML Calibration Register (BML/CAL/REG/001) | Entry made |  |  |

# 7.  Installation Qualification (IQ)

The IQ phase provides documented evidence that the pH meter has been supplied and installed in accordance with manufacturer's specifications, approved design documents, and BML requirements.

## 7.1  IQ Acceptance Criteria

All IQ tests must PASS before proceeding to OQ. Any FAIL requires documented deviation and corrective action.

## 7.2  IQ Test 1 – Document Verification

| Item | Expected / Acceptance Criterion | Result / Ref. |
| --- | --- | --- |
| User Requirements Specification (URS) | URS exists and is approved prior to purchase |  |
| Manufacturer's data sheet | Model, serial no., and specifications match PO |  |
| Factory calibration certificate | Certificate issued within 12 months; NIST-traceable |  |
| Operator's manual | Complete manual available in English |  |
| Packing list / invoice | All listed items received and accounted for |  |

## 7.3  IQ Test 2 – Physical and Hardware Verification

| Check | Acceptance Criterion | Observed | Pass/Fail |
| --- | --- | --- | --- |
| Main unit model | Matches PO and URS |  |  |
| Serial number | Serial no. matches CoC / invoice |  |  |
| Physical condition | No visible damage, cracks, or corrosion |  |  |
| Display screen | Fully legible, no dead pixels |  |  |
| Keypad / buttons | All functional; no sticking keys |  |  |
| Electrode socket | Clean, free of damage |  |  |
| ATC probe socket | Present and undamaged |  |  |
| Power adaptor / cable | Correct voltage rating; cable undamaged |  |  |
| Asset tag affixed | BML-EQ-PHM-001 tag visible |  |  |
| Installation location | Level surface, away from vibration and direct sunlight |  |  |
| Grounding | Unit properly earthed per manufacturer's instructions |  |  |

## 7.4  IQ Test 3 – Utilities Verification

| Utility | Acceptance Criterion | Observed | Pass/Fail |
| --- | --- | --- | --- |
| Mains voltage | Within manufacturer's specified range (e.g., 220–240 V AC ±10%) |  |  |
| Power socket | Dedicated, stable power supply; UPS preferred |  |  |
| Environmental temperature | 18–25°C at time of installation |  |  |
| Relative humidity | 30–70% RH |  |  |
| Bench area | Non-metallic, level surface with drainage proximity |  |  |

## 7.5  IQ Summary and Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 8.  Operational Qualification (OQ)

The OQ phase provides documented evidence that the pH meter operates consistently within defined limits throughout its intended operational range. All tests are performed by a trained analyst following BML/SOP/QC/005.

## 8.1  Buffer Standards Used

| Buffer | Nominal pH at 25°C | Lot / Expiry | Source | Traceability |
| --- | --- | --- | --- | --- |
| Buffer 1 | 4.00 ± 0.01 |  |  | NIST-traceable |
| Buffer 2 | 7.00 ± 0.01 |  |  | NIST-traceable |
| Buffer 3 | 10.01 ± 0.01 |  |  | NIST-traceable |

## 8.2  OQ Test 1 – Instrument Power-On and Self-Diagnostic

| Test | Acceptance Criterion | Result |
| --- | --- | --- |
| Power-on sequence | Instrument powers on without error codes |  |
| Self-diagnostic / RAM test | Internal self-test completes successfully |  |
| Display initialisation | All display segments visible; units displayed correctly |  |
| Date and time setting | Clock set accurately; retained after power cycle |  |
| Electrode connection recognition | Instrument recognises connected electrode without error |  |
| ATC probe response | ATC temperature reads ambient ± 0.5°C |  |

## 8.3  OQ Test 2 – Single-Point Calibration Verification (Neutrality)

Procedure: Rinse electrode with purified water, blot dry, immerse in pH 7.00 buffer. Allow reading to stabilise (≥1 min). Record displayed value. Acceptance: displayed pH = 7.00 ± 0.05.

| Replicate | Buffer pH | Displayed pH | Deviation | Pass/Fail |
| --- | --- | --- | --- | --- |
| 1 | 7.00 |  |  |  |
| 2 | 7.00 |  |  |  |
| 3 | 7.00 |  |  |  |

## 8.4  OQ Test 3 – Two-Point Calibration and Linearity (pH 4.00 and 10.01)

Procedure: Calibrate instrument using pH 4.00 buffer, then verify at pH 7.00, then pH 9.20. Record each displayed value. Acceptance criterion: ±0.05 pH units at each point.

| Buffer | Reference pH | Displayed pH | Deviation | Pass/Fail |
| --- | --- | --- | --- | --- |
| pH 4.00 buffer | 4.00 |  |  |  |
| pH 7.00 buffer | 7.00 |  |  |  |
| pH 10.01 buffer | 9.20 |  |  |  |

## 8.5  OQ Test 4 – Electrode Slope Check

The Nernst slope should be 95–105% of theoretical (59.16 mV/pH unit at 25°C). Most meters display slope% after two-point calibration. Acceptance: slope 95.0–105.0%.

| Parameter | Acceptance Criterion | Observed Value |
| --- | --- | --- |
| Electrode slope (%) | 95.0 – 105.0% |  |
| Asymmetry potential (mV) | ≤ ±30 mV from zero |  |

## 8.6  OQ Test 5 – Temperature Compensation Verification

Procedure: Prepare pH 7.00 buffer at two temperatures (approx. 20°C and 37°C) using a calibrated thermometer. Record displayed pH and temperature at each condition. Acceptance: ATC-corrected pH within ±0.05 of certified buffer value at each temperature.

| Buffer | Target Temp (°C) | Actual Temp (°C) | Expected pH | Displayed pH | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| pH 7.00 | 20 |  | 7.00 ± 0.05 |  |  |
| pH 7.00 | 37 |  | 7.00 ± 0.05 |  |  |

## 8.7  OQ Test 6 – Response Time

Procedure: Transfer electrode from pH 7.00 buffer to pH 4.00 buffer. Record time from immersion to stable reading (drift ≤ 0.01 pH/min). Acceptance: stable reading achieved within 60 seconds.

| Test | From Buffer | Time to Stable Reading (s) | Pass/Fail |
| --- | --- | --- | --- |
| Response time | pH 7.00 → pH 4.00 |  |  |
| Response time (reverse) | pH 4.00 → pH 10.01 |  |  |

## 8.8  OQ Test 7 – Reproducibility (Repeatability)

Procedure: Measure pH 7.00 buffer 10 consecutive times without re-calibrating. Record all values. Acceptance: RSD ≤ 0.3% and all readings within ±0.05 pH of reference.

| Reading No. | pH Displayed | Reading No. | pH Displayed | Criterion |
| --- | --- | --- | --- | --- |
| 1 |  | 6 |  |  |
| 2 |  | 7 |  | RSD ≤ 0.3% |
| 3 |  | 8 |  |  |
| 4 |  | 9 |  | All ± 0.05 pH |
| 5 |  | 10 |  |  |
| Mean ± SD: |  | Mean: _____ ± _____ SD; RSD: _____% |  |  |

## 8.9  OQ Summary and Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 9.  Performance Qualification (PQ)

The PQ phase provides documented evidence that the pH meter consistently performs within defined specifications when used for its intended purpose: measurement of pH in LVP pharmaceutical solutions. PQ is conducted under routine conditions including normal operator changeover.

## 9.1  PQ Test 1 – Inter-Day Precision (3 Days)

Procedure: Measure pH 7.00 buffer on three separate days (different analysts may be used). Calibrate fresh each day using pH 4.00 and 7.00 buffers. Acceptance: mean pH on each day within ±0.05; day-to-day RSD ≤ 0.5%.

| Day | Analyst | Buffer pH | Displayed pH | Deviation | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Day 1 |  | 7.00 |  |  |  |
| Day 2 |  | 7.00 |  |  |  |
| Day 3 |  | 7.00 |  |  |  |

## 9.2  PQ Test 2 – pH Measurement of LVP Products

Procedure: Obtain a representative sample of each LVP product (as approved batch or spiked reference). Measure pH using calibrated instrument. Compare against USP/Ph.Eur. specification. Conduct in triplicate on each product.

| Product | pH Specification (USP) | Result 1 | Result 2 | Result 3 | Pass/Fail |
| --- | --- | --- | --- | --- | --- |
| Normal Saline 0.9% (NaCl) | 4.5 – 7.0 |  |  |  |  |
| Dextrose 5% (D5W) | 3.2 – 6.5 |  |  |  |  |
| Ringer's Lactate | 5.0 – 7.0 |  |  |  |  |

## 9.3  PQ Test 3 – Accuracy Against External Reference

Procedure: A certified external reference pH measurement (NIST-traceable or by accredited external lab) is compared against BML instrument readings on the same sample. Acceptance: BML result within ±0.05 pH of certified reference.

| Sample | Reference pH (Certified) | BML Measured pH | Difference | Pass/Fail |
| --- | --- | --- | --- | --- |
| Sample A |  |  |  |  |
| Sample B |  |  |  |  |

## 9.4  PQ Test 4 – Operator-to-Operator Variability

Procedure: Two trained analysts independently measure the same pH 7.00 buffer sample (five readings each, same calibration). Acceptance: combined RSD ≤ 0.5%; mean difference between analysts ≤ 0.05 pH.

| Analyst A | pH Reading | Analyst B | pH Reading | Acceptance |
| --- | --- | --- | --- | --- |
| Reading 1 |  | Reading 1 |  | RSD ≤ 0.5% |
| Reading 2 |  | Reading 2 |  |  |
| Reading 3 |  | Reading 3 |  | Mean diff ≤ 0.05 pH |
| Reading 4 |  | Reading 4 |  |  |
| Reading 5 |  | Reading 5 |  |  |

## 9.5  PQ Summary and Sign-off

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 10.  Deviation Management

Any test result that fails to meet the acceptance criterion constitutes a deviation. The following procedure applies:

- The deviation shall be documented in the Qualification Report using the Deviation Log (Section 5 of the Report).

- An immediate investigation shall be conducted to determine root cause.

- Corrective and Preventive Action (CAPA) shall be assigned with a responsible person and target date.

- The failed test shall be repeated after CAPA implementation with documented justification.

- Major deviations affecting product safety or regulatory compliance shall be escalated to the QA Manager immediately.

- A requalification decision shall be made by the QA Manager based on the severity of the deviation.

No qualification phase shall be considered complete while open deviations with incomplete CAPA remain.

# 11.  Requalification Criteria

Full or partial requalification shall be triggered upon occurrence of any of the following events:

| Trigger Event | Qualification Phase | Frequency |
| --- | --- | --- |
| Routine periodic requalification | PQ (abbreviated) | Annual |
| Electrode replacement | OQ + PQ | As needed |
| Relocation to new bench or room | IQ + OQ + PQ | As needed |
| Major repair or firmware/software update | OQ + PQ | As needed |
| OOS trend identified during routine calibration | OQ + PQ | As needed |
| Failure identified during internal audit or inspection | Full IQ/OQ/PQ | As needed |

# 12.  Document Control

| Version | Date | Description of Change | Approved by |
| --- | --- | --- | --- |
| 01 | April 2025 | Initial issue | QA Manager |

# 13.  Appendix

## Appendix A – Buffer pH vs Temperature Reference Table

Reference values per NIST SRM 185j and USP <791>. ATC should correct for these shifts automatically.

| Temperature (°C) | pH 4.00 Buffer | pH 7.00 Buffer | pH 10.01 Buffer |
| --- | --- | --- | --- |
| 15 | 4.00 | 7.10 | 10.12 |
| 20 | 4.00 | 7.07 | 10.06 |
| 25 | 4.00 | 7.00 | 10.01 |
| 30 | 4.01 | 6.99 | 9.97 |
| 37 | 4.02 | 6.98 | 9.91 |

## Appendix B – Glossary

| Term | Definition |
| --- | --- |
| ATC | Automatic Temperature Compensation – electronic correction of pH reading for sample temperature |
| IQ | Installation Qualification – verification that equipment is installed correctly |
| Nernst Slope | Theoretical voltage change per pH unit at 25°C: 59.16 mV/pH. Electrode slope expressed as % of theoretical |
| NIST | National Institute of Standards and Technology (USA) – provides traceable reference standards |
| OQ | Operational Qualification – verification that equipment operates within specified ranges |
| PQ | Performance Qualification – verification that equipment consistently performs for its intended use |
| RSD | Relative Standard Deviation – SD / Mean × 100%; measure of precision |
| URS | User Requirements Specification – document defining what the instrument must do |