LOCATION – WTP/Lab

PROTOCOL No. – BML/QA/pHQ/001

Effective date: August, 2025.

Overall Status:

PROVIDED BY:

BIOMEDICAL LIMTED

KWARA STATE, NIGERIA

BIOMEDICAL LIMITED

pH METER QUALIFICATION REPORT

Installation Qualification / Operational Qualification / Performance Qualification

Contents

1.  Objective	4

2.  Equipment Identification	4

3.  Pre-Qualification Checklist Results	4

4.  Installation Qualification (IQ) Results	6

4.1  IQ Test 1 – Document Verification	6

4.2  IQ Test 2 – Physical and Hardware Verification	6

4.3  IQ Test 3 – Utilities Verification	7

4.4  IQ Conclusion	7

5.  Operational Qualification (OQ) Results	8

5.1  OQ Test 1 – Power-On and Self-Diagnostic Results	8

5.2  OQ Test 2 – Single-Point Calibration Verification	8

5.3  OQ Test 3 – Two-Point Calibration and Linearity	9

5.4  OQ Test 4 – Electrode Slope Check	9

5.5  OQ Test 5 – Temperature Compensation Verification	9

5.6  OQ Test 6 – Response Time	9

5.7  OQ Test 7 – Repeatability (10 Consecutive Readings)	10

5.8  OQ Conclusion	10

6.  Performance Qualification (PQ) Results	11

6.1  PQ Test 1 – Inter-Day Precision (3 Days)	11

6.2  PQ Test 2 – pH Measurement of LVP Products	11

6.3  PQ Test 3 – Accuracy Against External Reference	11

6.4  PQ Test 4 – Operator-to-Operator Variability	12

6.5  PQ Conclusion	12

7.  Deviation Log	13

8.  Overall Qualification Summary and Release Decision	13

8.1  Final Qualification Conclusion	13

8.2  Final Release Signatures	14

9.  Post-Qualification Requirements and Routine Maintenance	14

10.  Document Control	15

11.  Appendix – Raw Data Attachment Sheet	15

REPORT FORMAT APPROVAL:

Signatures on this page indicate that this document and its format have been prepared, checked and approved by the personnel concerned. It shall be signed and date as shown below

|  | PREPARED BY: | REVIEWED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

|  | APPROVED BY: | AUTHORIZED BY: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

# 1.  Objective

This Qualification Report documents the results, conclusions, and approval status of the IQ, OQ, and PQ activities performed on the pH meter installed in the Plant at Biomedical Ltd (BML). The qualification was executed in accordance with Protocol BML/QA/pHQ/001 Version 01.

This report provides the formal documented evidence required by EU GMP Annex 15, USP <791>, and NAFDAC GMP Guidelines that the pH meter is fit for its intended purpose and approved for routine use in pharmaceutical manufacturing quality control.

# 2.  Equipment Identification

| Parameter | Details (to be completed at time of execution) |
| --- | --- |
| Equipment Name | pH Meter – Laboratory/Process Grade |
| Manufacturer | ___________________________________ |
| Model | ___________________________________ |
| Serial Number | ___________________________________ |
| Electrode Type / Serial No. | ___________________________________ |
| ATC Probe Serial No. | ___________________________________ |
| Factory Calibration Cert. No. | ___________________________________ |
| Location | QC Laboratory / WTP |
| Qualification Start Date | ___________________________________ |
| Qualification End Date | ___________________________________ |

# 3.  Pre-Qualification Checklist Results

| Pre-Qualification Check | Criterion | Status | Init. |
| --- | --- | --- | --- |
| Purchase Order and delivery documentation available | Documents on file |  |  |
| Instrument received undamaged – visual inspection | No damage |  |  |
| Manufacturer's CoC / Factory Calibration Certificate received | CoC on file |  |  |
| Approved operating SOP (BML/SOP/QC/005) in place | SOP current |  |  |
| NIST-traceable buffer standards procured (pH 4.00, 7.00, 9.20) | Standards available |  |  |
| Environmental conditions within specification | 18–25°C / 30–70% RH |  |  |
| Analyst trained on SOP | Training record |  |  |
| Instrument registered in Calibration Register BML/QA/CR/001 | Entry made |  |  |

Note: 'PENDING' cells to be updated to PASS/FAIL upon execution. All items must PASS before IQ commences.

# 4.  Installation Qualification (IQ) Results

## 4.1  IQ Test 1 – Document Verification

| Document | Acceptance Criterion | Reference / Finding | Status |
| --- | --- | --- | --- |
| URS | Approved URS on file prior to purchase |  |  |
| Manufacturer's data sheet | Matches PO specifications |  |  |
| Factory calibration certificate | NIST-traceable; within 12 months |  |  |
| Operator's manual | Complete manual in English |  |  |
| Packing list / invoice | All items received and accounted for |  |  |

## 4.2  IQ Test 2 – Physical and Hardware Verification

| Check | Acceptance Criterion | Observed | Status |
| --- | --- | --- | --- |
| Main unit model | Matches PO and URS |  |  |
| Serial number | Matches CoC / invoice |  |  |
| Physical condition | No damage, cracks, or corrosion |  |  |
| Display screen | Fully legible, no dead pixels |  |  |
| Keypad / buttons | All functional; no sticking keys |  |  |
| Electrode socket | Clean, undamaged |  |  |
| ATC probe socket | Present and undamaged |  |  |
| Power adaptor / cable | Correct voltage; cable undamaged |  |  |
| BML asset tag affixed | BL/EQ/QAE/ visible |  |  |
| Installation location | Level; away from vibration and direct sunlight |  |  |
| Grounding | Properly earthed per manufacturer |  |  |

## 4.3  IQ Test 3 – Utilities Verification

| Utility | Acceptance Criterion | Observed | Status |
| --- | --- | --- | --- |
| Mains voltage | Within manufacturer's range (e.g., 220–240 V AC ±10%) |  |  |
| Power socket | Stable; UPS preferred |  |  |
| Environmental temperature | 18–25°C |  |  |
| Relative humidity | 30–70% RH |  |  |
| Bench area | Non-metallic, level, with drainage proximity |  |  |

## 4.4  IQ Conclusion

All IQ tests are pre-populated as PENDING. Upon execution, each test outcome shall be recorded in the above tables and the following conclusion completed.

| IQ Outcome | Details |
| --- | --- |
| Overall IQ Status |  |
| Number of deviations |  |
| Deviation references |  |
| IQ Conclusion statement | The pH meter has been / has not been [delete as applicable] successfully installed in accordance with Protocol BML/QA/pHQ/001. |

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 5.  Operational Qualification (OQ) Results

Buffer standards used and traceability details to be recorded in the table below at time of execution.

| Buffer | Nominal pH at 25°C | Lot / Expiry | Source | Traceability |
| --- | --- | --- | --- | --- |
| Buffer 1 | 4.00 ± 0.01 |  |  | NIST-traceable |
| Buffer 2 | 7.00 ± 0.01 |  |  | NIST-traceable |
| Buffer 3 | 9.20 ± 0.01 |  |  | NIST-traceable |

## 5.1  OQ Test 1 – Power-On and Self-Diagnostic Results

| Test | Acceptance Criterion | Result | Status |
| --- | --- | --- | --- |
| Power-on sequence | Powers on without error codes |  |  |
| Self-diagnostic / RAM test | Self-test completes successfully |  |  |
| Display initialisation | All segments visible; units displayed correctly |  |  |
| Date and time setting | Clock accurate; retained after power cycle |  |  |
| Electrode connection recognition | Recognised without error |  |  |
| ATC probe response | ATC temperature = ambient ± 0.5°C |  |  |

## 5.2  OQ Test 2 – Single-Point Calibration Verification

| Replicate | Buffer pH | Displayed pH | Deviation | Status |
| --- | --- | --- | --- | --- |
| 1 | 7.00 |  |  |  |
| 2 | 7.00 |  |  |  |
| 3 | 7.00 |  |  |  |
| Acceptance | ±0.05 pH | Mean: |  |  |

## 5.3  OQ Test 3 – Two-Point Calibration and Linearity

| Buffer | Reference pH | Displayed pH | Deviation | Status |
| --- | --- | --- | --- | --- |
| pH 4.00 buffer | 4.00 |  |  |  |
| pH 7.00 buffer | 7.00 |  |  |  |
| pH 9.20 buffer | 9.20 |  |  |  |

Acceptance: ±0.05 pH units at each point.

## 5.4  OQ Test 4 – Electrode Slope Check

| Parameter | Acceptance Criterion | Observed | Status |
| --- | --- | --- | --- |
| Electrode slope (%) | 95.0 – 105.0% |  |  |
| Asymmetry potential (mV) | ≤ ±30 mV |  |  |

Note: A slope below 95% indicates electrode degradation. Replace electrode and re-test before proceeding.

## 5.5  OQ Test 5 – Temperature Compensation Verification

| Buffer | Target Temp (°C) | Actual Temp (°C) | Expected pH | Displayed pH | Status |
| --- | --- | --- | --- | --- | --- |
| pH 7.00 | 20 |  | 7.00 ± 0.05 |  |  |
| pH 7.00 | 37 |  | 7.00 ± 0.05 |  |  |

## 5.6  OQ Test 6 – Response Time

| Test | From Buffer | Time to Stable Reading (s) | Criterion | Status |
| --- | --- | --- | --- | --- |
| Response time | pH 7.00 → pH 4.00 |  | ≤ 60 s |  |
| Response time (reverse) | pH 4.00 → pH 9.20 |  | ≤ 60 s |  |

## 5.7  OQ Test 7 – Repeatability (10 Consecutive Readings)

| Reading | pH Displayed | Reading | pH Displayed | Criterion |
| --- | --- | --- | --- | --- |
| 1 |  | 6 |  |  |
| 2 |  | 7 |  | RSD ≤ 0.3% |
| 3 |  | 8 |  |  |
| 4 |  | 9 |  | All ± 0.05 pH |
| 5 |  | 10 |  |  |
| Mean ± SD |  | Mean: ________ SD: ________ RSD: ________% |  | Status: |

## 5.8  OQ Conclusion

| OQ Outcome | Details |
| --- | --- |
| Overall OQ Status | PENDING – complete upon execution |
| Number of deviations |  |
| Deviation references |  |
| OQ Conclusion statement | The pH meter has been / has not been demonstrated to operate within specified limits as required by Protocol BML/QA/pHQ/001. |

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 6.  Performance Qualification (PQ) Results

## 6.1  PQ Test 1 – Inter-Day Precision (3 Days)

| Day | Analyst | Buffer pH | Displayed pH | Deviation | Criterion | Status |
| --- | --- | --- | --- | --- | --- | --- |
| Day 1 |  | 7.00 |  |  | ±0.05 / RSD ≤0.5% |  |
| Day 2 |  | 7.00 |  |  | ±0.05 / RSD ≤0.5% |  |
| Day 3 |  | 7.00 |  |  | ±0.05 / RSD ≤0.5% |  |

## 6.2  PQ Test 2 – pH Measurement of LVP Products

All measurements to be performed in triplicate on representative samples from approved batches or reference samples. Results compared against USP/Ph.Eur. specifications.

| Product | Specification (USP) | Result 1 | Result 2 | Result 3 | Mean | Status |
| --- | --- | --- | --- | --- | --- | --- |
| Normal Saline 0.9% | 4.5 – 7.0 |  |  |  |  |  |
| Dextrose 5% (D5W) | 3.5 – 6.5 |  |  |  |  |  |
| Ringer's Lactate | 6.0 – 7.5 |  |  |  |  |  |

## 6.3  PQ Test 3 – Accuracy Against External Reference

| Sample | Reference pH (Certified) | BML Measured pH | Difference | Status |
| --- | --- | --- | --- | --- |
| Sample A |  |  |  |  |
| Sample B |  |  |  |  |

Acceptance: BML result within ±0.05 pH of certified reference.

## 6.4  PQ Test 4 – Operator-to-Operator Variability

| Analyst A | pH Reading | Analyst B | pH Reading | Statistical Summary |
| --- | --- | --- | --- | --- |
| Reading 1 |  | Reading 1 |  | Analyst A Mean: |
| Reading 2 |  | Reading 2 |  | Analyst B Mean: |
| Reading 3 |  | Reading 3 |  | Combined RSD: |
| Reading 4 |  | Reading 4 |  | Mean Difference: |
| Reading 5 |  | Reading 5 |  | Overall Status: |

Acceptance: Combined RSD ≤ 0.5%; mean difference between analysts ≤ 0.05 pH.

## 6.5  PQ Conclusion

| PQ Outcome | Details |
| --- | --- |
| Overall PQ Status |  |
| Number of deviations |  |
| Deviation references |  |
| PQ Conclusion statement | The pH meter has been / has not been demonstrated to perform consistently for its intended use in LVP pharmaceutical manufacturing. |

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| Executed by (Analyst) |  |  |  |
| Reviewed by (QAQC Head) |  |  |  |
| Approved by (QA Manager) |  |  |  |

# 7.  Deviation Log

All deviations observed during qualification activities shall be recorded in this log. No deviations identified = enter 'None'.

| Dev. No. | Test Ref. | Date | Description of Deviation | CAPA / Corrective Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 |  |  | None | N/A | Closed |

# 8.  Overall Qualification Summary and Release Decision

| Qualification Phase | Scope | # Tests | Overall Status |
| --- | --- | --- | --- |
| Pre-Qualification Checks | 8 administrative and environmental checks | 8 |  |
| Installation Qualification (IQ) | Document verification, physical checks, utilities | 3 test groups |  |
| Operational Qualification (OQ) | Power-on, calibration, linearity, slope, ATC, response time, repeatability | 7 tests |  |
| Performance Qualification (PQ) | Inter-day precision, product testing, external accuracy, operator variability | 4 tests |  |

## 8.1  Final Qualification Conclusion

CONCLUSION (to be completed upon execution): Based on the qualification activities performed in accordance with Protocol BML/QA/pHQ/001 Version 01, the pH meter (Equipment Tag: ) has been / has not been demonstrated to be: - Correctly installed in accordance with manufacturer's specifications and BML requirements (IQ PASSED / FAILED) - Operating within defined performance limits across its intended measurement range (OQ PASSED / FAILED) - Consistently performing accurate pH measurements suitable for LVP pharmaceutical manufacturing QC (PQ PASSED / FAILED) The pH meter is therefore APPROVED / NOT APPROVED for routine use. Routine calibration shall be performed daily (before use) in accordance with BML/QDG/QA/ , with quarterly external calibration as per BML/QA/CAL/001.

## 8.2  Final Release Signatures

| Role | Name | Signature | Date |
| --- | --- | --- | --- |
| QA officer |  |  |  |
| QA Manager |  |  |  |
| Factory Manager |  |  |  |

# 9.  Post-Qualification Requirements and Routine Maintenance

| Activity | Details | Frequency | SOP Ref. |
| --- | --- | --- | --- |
| Routine calibration (2-point) | pH 4.00 and pH 7.00 buffers; before each use | Daily / Per use |  |
| Electrode conditioning | Store electrode in KCl solution; condition in pH 7.00 buffer before use | Daily |  |
| Slope check | Verify slope 95–105% after 2-point calibration | Daily |  |
| Electrode visual inspection | Check for cracks, contamination, or electrolyte depletion | Weekly |  |
| External calibration | By accredited NIST-traceable calibration laboratory | Annual |  |
| Requalification review | Abbreviated PQ using 3 LVP products; review calibration trends | Annual | This report |
| Electrode replacement | Replace when slope <95% or asymmetry >30 mV; trigger OQ+PQ | As needed |  |

# 10.  Document Control

| Version | Date | Description of Change | Approved by |
| --- | --- | --- | --- |
| 01 | April 2026 | Initial issue – qualification report template | QA Manager |

# 11.  Appendix – Raw Data Attachment Sheet

The following raw data worksheets shall be attached to this report upon completion of qualification activities:

| # | Attachment Description | Document Reference | Attached (Y/N) |
| --- | --- | --- | --- |
| A | Completed IQ data sheets with analyst initials and dates | Section 7 |  |
| B | Completed OQ data sheets (all 7 tests) | Section 8 |  |
| C | Completed PQ data sheets (all 4 tests) | Section 9 |  |
| D | Manufacturer's factory calibration certificate | CoC No.: ___________ |  |
| E | Buffer standard certificates of analysis | Lot No.(s): ___________ |  |
| F | Analyst training records for all personnel who executed tests | BML Training System |  |
| G | Deviation forms (if any) | DEV/ |  |
| H | Environmental monitoring records at time of qualification | Temp/RH log |  |