- OBJECTIVE

To define the procedure and schedule for the calibration and maintenance of equipment, ensuring reliable and accurate results for all processes and testing activities that may affect product quality.

- SCOPE

This SOP is applicable to all critical and non-critical instruments, equipment, and measuring devices used in the facility, including those in manufacturing, quality control (QC), engineering, and warehouse departments

- RESPONSIBILITY

- User Departments (QC, Production, Warehouse, etc.): Responsible for ensuring their instruments are included in the calibration program, checking the calibration status before use, and initiating a deviation if a due date is missed or an issue is found.

- Engineering /QC Department: Responsible for preparing the annual and monthly calibration schedules, performing calibrations (in-house or coordinating external agencies), and maintaining all calibration records.

- Quality Assurance (QA) Department: Responsible for reviewing and approving calibration SOPs, schedules, and certificates, ensuring overall compliance, and evaluating the impact of any out-of-calibration (OOC) situations on product quality.

- ABBREVIATION AND DEFINITIONS

- Calibration: is the process of comparing reading on one piece of equipment or system with another piece of equipment that has been calibrated and referenced to a known set of parameters.

- Accuracy: the closeness of agreement between an observed or measured value and accepted reference value.

- Measurement standards: a measuring instrument or material, which physically defines a unit of measurement or value of a quantity. Measurement standards should be traceable to a national or international standard.

- Critical instruments: the instrument, which gives a direct impact on the product.

- Non-critical instrument: the instrument which does not have a direct impact on the product and is optional

- NIST: National Institute of Standards and Technology

- SON: Standard organization of Nigeria

- CAC: Corporate Affairs Commission

- IPAN: Institute of Public Analysts of Nigeria

- PROCEDURES

- According to the calibration policy, all master instruments must be calibrated at once in a year or as desired in case of special requirement and shall be traceable to national standards like IPAN, SON, NIST.

- All measuring and monitoring devises/instruments are calibrated either internally or by an approved external calibrator.

Classification of Instruments:

- All calibrated instruments will be tagged with an identification number for tracking purposes as per Calibration status label. This tag will also indicate the date of last calibration and the due date of the next calibration.

Equipment Calibration Status label

CALIBRATION STATUS

EQUIPMENT/INSTRUMENT NAME:

IDENTIFICATION CODE:

MODEL                                        FREQUENCY

DONE ON:                                   DUE ON:

STATUS:			      DATE OF LAST CALIBRATION:

DONE BY:                                   CHECKED BY:

DATE:                                           DATE:

- Instruments deemed not to require calibration due to the nature of the instrument or its application will be tagged as “NO CALIBRATION REQUIRED”.

- When any instrument is out of calibration limits, it should be immediately labeled as “OUT OF CALIBRATION” as per annexure 6 and intimated to Head -Engineering and get the same recalibrated, repaired, replaced as the case may be before proceeding further.

Out of calibration label

OUT OF CALIBRATION

SIGN:

DATE:

DO NOT USE

- The intimation of the same shall be given to Head QA to assess the impact of such out of calibration instrument on product quality, if instrument has been used, and during this period out of calibration has been observed.

Calibration Intervals

- Calibration Intervals will be assigned to each type of instrument to attain a desired probability that the instrument will be within calibration limits at the time of use.

- All instruments will be assigned calibration frequency as mentioned in below table.

| Sr. No. | Calibration / Validation Frequency | ± Limit |
| --- | --- | --- |
| 01 | 1 Month (Monthly-Once in a month) | ± 3 working days |
| 02 | 3 Months (Quarterly-Once in three month) | ± 7 working days |
| 03 | 6 Months (Half yearly-Once in six month) | ± 15 working days |
| 04 | 12 Months (Yearly-Once in year) | ± 30 working days |

- Initial interval of calibration for new equipment will be based on the manufacturer’s recommendation and judgment of engineering department.

- Calibration intervals may be adjusted as necessary on the basis of calibration data or other information which may support for change. These changes can be done as per procedure of Change control.

- The date of calibration shall be put on the calibration label/ documents however the next due date of calibration will be the one day before on which calibration can be performed, for example current date of calibration is 10 Jan 2022 then the due date of calibration for 12 months period shall be 09 Jan 2023.

- Maintain the Master List of Calibration Status of Measuring Devices of Equipment/ Instruments/Utilities as per Enclosure –1 in soft copy and hard copy.

- When the need for repair/ adjustment is determined, the instrument repaired or adjusted by in-house or outside contractor will be re-calibrated prior to returning to service.

- When the need for repair/ adjustment is determined, the instrument will be tagged as “OUT OF CALIBRATION”.

- The “OUT OF CALIBRATION” tag will not be removed until the repair is completed and the instrument has been re-calibrated.

Traceability

- The calibration system will provide the means to trace assigned values of measurement to nationally recognized standards (i.e. NABL). Documentation will include calibration agency and certification data.

- The use of calibration standards which are not traceable to National standards, such as standard reference materials or internal standards will require adequate documentation as to the validity of the standards used.

Outside Contractor Control for External calibration

- It is the responsibility of company Engineering Department in co-ordination with QA to ensure that the contractor’s calibration system complies with all requirements;

  - Requirement for External Calibration Service

  - CAC registration

  - ISO/IEC 17025 accreditation certificate (or equivalent evidence of traceability such as SON/IPAN certificate etc.)

  - Sample calibration certificate (for the instrument type in a pharmaceutical Laboratory).

  - Calibration Equipment Certificate

  - List of reference standards with last calibration dates

  - Technician CVs and training records, and two client references (preferably pharmaceutical).

  - Organization structure

  - Vendor qualification will be conducted using appropriate documents

- The Head QA and Head Engineering Department will review the outside contractor’s procedure and will approve.

- All documentation supplied by contractor will be reviewed and dully signed.

- Calibration of instruments/equipment/measuring devices shall be done as per standard operating procedures of respective instruments.

- Calibration Process Limits shall be defined for each equipment /instrument/ measuring device.

- There shall be a means of readily determining the calibration status of each equipment/instrument/ measuring devices. The instrument/equipment measuring devices shall be properly labeled with “calibration status label” as described in respective standard operating procedures of calibration. Wherever the calibration of equipment/instrument/measuring devices is coordinated by outside calibration agency then calibration status label will be supplied by outside calibration agency.

Calibration Certificate should contain following details:

  - Instrument/equipment/measuring device name

  - Instrument/equipment/measuring device Identification Number

  - Range of instruments under calibration

  - Name of Standard Instrument and its range used in calibration process

  - Calibration Certificate of standard Instrument used in calibration process and validity period of calibration

  - Date and time of calibration

  - Calibration status

  - Calibration due date

  - Acceptance criteria

- Master List of Calibration Status of Measuring Devices of Equipment/ Instruments/Utilities shall prepare separately and updated by engineering department. This document shall contain name and identification number of instrument /equipment/ measuring device, location of instrument/equipment/measuring device, date of calibration, due date of calibration.

Note: Measuring devices which are in spare shall be calibrated as per calibration frequency. These calibrated measuring devices shall be used as replacement in case of urgency like mal-functioning of measuring devices or if calibration validity of measuring devices provided to equipment /systems/areas is immediately due.

- All QC instruments shall calibrate as per standard calibration procedure of respective instrument.

- Any out of calibration instrument/measuring devices shall labeled as ‘OUT OF CALIBRATION’ and shall intimate to HOD of concern department, Head Engineering and Head QA. Such instruments/measuring devices shall not use till the satisfactory recalibration of the same. In case of critical instruments which have direct impact on product quality / analysis results found out of calibration, the products/samples which analyzed by using instrument within the period of previous calibration date and date on which instrument found out of calibration shall reassessed.

- Instruments calibrated internally includes; Glass thermometer, micrometer, digital veneer caliper, stop watch, pipette, burette, glass wares.

REFERENCE AND ENCLOSURE

Enclosure 1 - Annual Calibration Schedule form- BML-QDG-QA-031-F01

Enclosure 2- Instrument / Equipment Master List - BML-QDG-QA-031-F02

Enclosure 3 – internal calibration schedule- BML-QDG-QA-031-F03

Enclosure 4 – External calibration schedule- BML-QDG-QA-031-F04

7.0 TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff | All staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |