- OBJECTIVE:

This SOP is prepared to have a laid down standard Procedure for Calibration of UV/ Visible Spectrophotometer in the Chemical Laboratory.

- SCOPE:

To provide written procedure for calibration of UV Visible Spectrophotometer.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Quality Control Analyst is responsible to calibrate the instrument as per SOP.

- Head QC: Ensure the compliance of SOP.

- QA Head shall be Accountable for implementation of SOP.

- ABBREVIATION AND DEFINITIONS

No.: Number

UV: Ultraviolet

KOH: Potassium Hydroxide

HPLC: High Performance Liquid Chromatography

SOP: Standard Operating Procedure

QC: Quality Control

- PROCEDURES

A UV-VIS spectrophotometer plays a crucial role in analyzing and characterizing various samples by measuring their interaction with light in the ultraviolet (UV) and visible (VIS) regions of the electromagnetic spectrum

- All calibration standards used for the calibration should be prepared as per SOP.

- Only use calibration standard which is procured from authentic source with certification and it should be pharmacopeia reference standards or from instrument manufacturer.

INTERNAL CALIBRATION:

Initial Set-up

- Check the rear panel Voltage Selector switch is set to correct position for the power supply being used.

- Connect to the correct voltage setting & switch on the power supply.

- The Spectrophotometer will check the position of its light sources, mirrors and second order suppression filters after every power on cycle. It will take some time & also lamp will take some time to stabilize. So allow the machine to stabilize for 30 minutes ( Half early)

Cuvette Qualification

- Clean the cell/cuvette with methanol

- Remove any spot if any

- Then rinse and fill with distilled water.

- Measure its apparent absorbance against air at 240nm for quartz cells and 650nm for glass cells.

- The apparent absorbance should not be greater than 0.093 for 1cm quartz cells (UV region) and 0.035 for 1cm glass cells (visible region).

- After that measure, the apparent absorbance after rotate the cell/cuvette in its holder (1800) again.

- Check the absorbance, rotating the cells should give the absorbance difference not greater than 0.005 from initial.

- Record the observations.

Control of Absorbance:

- Dry a quantity of Potassium dichromate (AR Grade) by heating to constant weight at 130°C.

- Take 100mg dry Potassium dichromate in 100ml vol. flask and allow to dissolve in 0.01N H2SO4 dilute up to the mark with 0.01N H2SO4 solution.

- Take 6ml of resulting solution dilute to 100ml with 0.01N H2SO4.

- Measure the absorbance of the solution at 235nm, 257nm, 313nm & 350nm taking 0.01N H2SO4 solution as blank.

Control of wavelengths:

- Prepare the solution of holmium per chlorate as follows.

- Dissolve 0.5 gm of holmium oxide in 2.4 ml of perchloric acid AR by warming gently and diluting to 10ml with purified water.

- Check the extinction of the solution, permitted tolerance is 1nm for the range 200 to 400 nm and 3 nm for the range 400 to 600 nm.

Limit of Stray Light:

- Prepare 1.2% w/v solution of Potassium Chloride.

- Measure the absorbance at 200nm.

- Absorbance should be not less than 2.0.

Resolution Power:

- Prepare a 0.02% v/v solution of Toluene (HPLC grade) in Hexane (HPLC grade).

- Record the ratio of the absorbance at the maxima at about 269 nm to that at the minima at about 266 nm.

- Ratio should not be not less than 1.5.

- Report the result in the format.

Frequency: Half early i.e 30 minutes before the schedule time

Action: If the calibration is not proper then repeat the procedure and report the results to the department head for an appropriate action.

- REFERENCE AND ENCLOSURE

- SOP Training on manual

- Enclosure 1- Calibration Record of UV Spectrophotometer- BML-QDG-QA- FO1

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | Quality Control Analysts |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |