|  | Biomedical Limited, No. 1 Ohimege Road, Industrial Estate, Ilorin, Kwara State, PMB 1449 | Doc. No.: BML/QDG/QA/038 |
| --- | --- | --- |
|  | QUALITY ASSURANCE DEPARTMENT EQUIPMENT CALIBRATION LOG UV-SPECTROPHOTOMETER | Version No.: 02 |

Calibration Date:

Next due date:

- Control  of Absorbance

| Wavelength (nm) | Absorbance (1%, 1cm) | Absorbance observed (1%, 1cm) | Tolerance |
| --- | --- | --- | --- |
| 235 | 124.5 |  | 122.9 – 126.2 |
| 257 | 144.5 |  | 142.8 – 148.2 |
| 313 | 48.6 |  | 47.0 – 50.3 |
| 350 | 107.3 |  | 105.6 – 109.0 |

Result: Pass/Fail

- Control of Wavelength.

| Wavelength at Absorption Maxima |  | Tolerance limit (nm) |
| --- | --- | --- |
| Expected wavelength (nm) | Observed wavelength (nm) |  |
| 241.15 |  | 240.15 – 242.15 |
| 287.15 |  | 286.15 – 288.15 |
| 360.50 |  | 360. 50 – 362.50 |
| 536.30 |  | 533.30 – 539.30 |

Result: Pass/Fail

Analysed by:                                                                                 Checked by:

Date :                                                                                              Date :

- Limit of Stray Light

__________________gm (1.2gm of KCl) into 100ml with purified water. Measure the absorbance at 200nm with water as blank.

| Wavelength | Absorbance | Tolerance limit |
| --- | --- | --- |
|  |  | Should be greater than 2.0 at a wavelength between 198nm to 202nm |

Result: Pass/Fail

- Resolution power

Record the spectrum of a 0.02%v/v solution of Toluene (HPLC grade) in Hexane (HPLC grade)

- Absorbance at the maxima at about 269

- Absorbance at the minima at about 266

Limit NLT 1.5

________________________________

Q.C Head

Sign /Date

EQUIPMENT CALIBRATION LOG BOOK

| Equipment Name: | UV-SPECTROPHOTOMETER |
| --- | --- |
| Equipment ID: | BL/EQ/QAE/001 |
| Location: | CHEMICAL |
| Model: | 752N |
| Manufacturer’s Name: | EQUIPTRONICS, INDIA |
| Issued By | QA DOCUMENTATION UNIT |

BIOMEDICAL LTD.

BIOMEDICAL LIMITED, NO 1 OHIMEGE ROAD, INDUSTRIAL ESTATE, ILORIN KWARA STATE, PMB 1449

QUALITY CONTROL

[CHEMICAL LAB]

CALIBRATION LOGBOOK FOR UV-SPECTROPHOTOMETER

LGB/QDG/QA/03

VOL. 02