- Objective: To lay down the procedure for calibration and temperature mapping of Refrigerator in the Microbiological Laboratory

- Scope: This SOP is applicable for the procedure for calibration and temperature mapping of deep freezer at Biomedical Limited.

- RESPONSIBILITY

- QC Officers / Analysts are responsible to perform the activity as per SOP

- Overall responsibility for training, implementation and follow-up with the QC Manager or his/her nominee

- ABBREVIATION AND DEFINITIONS

ID- Identification

No. Number

OC- Temperature degree Celsius

+ - Plus

- Minus

- PROCEDURES

- Place the calibrated Thermometer inside the Refrigerator and wait for 5 minutes.

- Observe the temperature for at least 3 hours and the temperature shall be recorded at an interval of 30 minutes.

- Note down the temperature of Thermometer and compare it with that of Refrigerator display temperature.

- The frequency of calibration shall be performed quarterly

- Maintain a Calibration Record for Refrigerator to note down the calibration result.

Temperature mapping procedure for Refrigerator

- Ensure that the preventive maintenance activity prior to temperature mapping of the refrigerator (Performance qualification).

- Temperature mapping shall be carried out for a period of 24 hours and the temperature shall be recorded at an interval of 15 minutes.

- Temperature mapping shall be carried out with six probes and the pattern

- Select upper, and lower racks in the refrigerator. Ensure that the calibration status of each probe before starting the calibration.

- Insert the probes (06 no’s) carefully in upper, middle and lower racks without disturbing the cultures/ samples.

- Temperature mapping shall be carried once in year by placing the probes in the upper, middle and lower racks, when in operation as a part of instrument performance check.

- The maximum and minimum reading shall be recorded from the raw data in the Log book.

- Channels shall be labeled to each probe.

- A variation of ± 5 °C is accepted between calibrated thermometer / probe and Refrigerator display temperature.

- If any discrepancy is observed during calibration, inform to the In-charge Quality Control and Maintenance Department for further corrective action.

TEMPERATURE:      0C (LIMIT +5%)

| S/N | TEMP. PROBE LOCATION | Observation Temp OC |  | Difference in OC | Pass/Fail |
| --- | --- | --- | --- | --- | --- |

- Frequency: once in a year

| Temperature probe ID No. | Calibrated on |  |
| --- | --- | --- |

Done by:                                                                                                                           Checked by:

Sign/Date                                                                                                                          Sign/Date

| Set Temperature | Tolerance (±2°C) | Temperature readings (°C) |
| --- | --- | --- |

UPPER

MIDDLE

LOWER

Temperature Mapping pattern-

PROBE 1	PROBE 2

PROBE 3	PROBE 4

PROBE 5

PROBE 7

PROBE 6

| Probe No. | Location |
| --- | --- |
| 1 | Left side of the lower rack upper tray |
| 2 | Right side of the lower rack upper tray |
| 3 | Left side of the lower rack middle tray |
| 4 | Right side of the lower rack middle tray |
| 5 | Centre of middle rack |
| 6 | Left side of the lower rack lower tray |
| 7. | Right side of the lower rack lower tray |

- REFERENCE AND ENCLOSURE

NIL

7.0 TRAINING INVOLVEMENT:

| Quality Control Manager | QC Head |
| --- | --- |
| QC Microbiologist | Lab. Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | MICROBIOLOGY LABORATORY |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |