- OBJECTIVE:

This SOP describes the procedure to verify the efficacy of steam sterilization cycles by challenging the autoclave with highly resistant microorganisms.

- SCOPE:

To provide written guidelines for establishing the efficacy of the Autoclave using (ATTEST(R)) Biological Indicators.

- RESPONSIBILITIES/ACCOUNTABILITIES:

Microbiologist, Quality Assurance Officer, Q.A Manager, Superintendent Pharmacist/Factory Manager.

- ABBREVIATION AND DEFINITIONS

Nil

- PROCEDURES

Take two vials of Attest Biological Indicator. Insert the vial into the crusher and crush properly.  Attach each vial of the Attest Indicator to a bag of infusion meant to be sterilized using paper tape.  Place one vial on the top tray and the other to the lower tray.  Allow the product to be autoclaved at 110oC, 1.7 bars and 60 minutes.  After the cycle, remove the vials and incubator in the Attest Incubator for a period of 7-14 days.

Incubation:

- Clean the inside of the stainless-steel incubator with 70% ethanol

- Allow to dry properly

- Pour 200ml of distilled water into the stainless-steel bowl of the incubator

- Insert the autoclaved vials into the plastic rack provided for it in the incubator.

- Cover the bowl with lid.

- Connect the incubator to power source, to enable it attain and maintain 60oC, a temperature required to stimulate the spores of Bacillus stearothermophilus to grow.

- Observe for 7-14 days.

Observation:

If the Spores of the Bacillus Stearothermophilus remained viable, the yellow strip in the vial shows creamy white or yellowish patch.  This indicates that the sterilization has not been effective by the autoclaving and there is growth.  If the strips retain the initial uniform colour, the sterilization is effective.

Results:

|  | Day 1 | Day 2 | Day 3 | Day 4 | Day 5 | Day 6 | Day 7 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Autoclaved Attest Vial | -ve | -ve | -ve | -ve | -ve | -ve | -ve |
| Control | +ve | +ve | +ve | +ve | +ve | +ve | +ve |

+ve	-	growth

-ve	-	no growth (sterile)

Sterilization is effective and hence Autoclaving Validated if the strips in the autoclaved vials retain the initial uniform colour, without any whitish or yellow patch (-ve)

Control:

Two un-autoclaved vials of Attest 3M are crushed and incubated using Attest incubator.  They grow within 24-48 hours and show creamy white or yellowish patches on the strip inside the vials +ve.

Conclusion:

Autoclaving is validated when the process can suitably render a viable heat resistant strain of Bacillus stearothermophilus not viable.]

Materials

- Attest® 3M incubator

- Distilled water

- Attest 3M biological indicator (containing Bacillus Stearothermophilus spores)

- Autoclave

- Attest indicator crusher.

- REFERENCE AND ENCLOSURE

- In-house Method

7.0 TRAINING INVOLVEMENT

| Quality Control Manager | Quality Control Microbiologist |
| --- | --- |
| QA staff | Laboratory Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in SOP format. - Inclusion of traceability of enclosures to parent SOP |