- OBJECTIVE:

To provide written guidelines for Chemical Analysis of water.

- SCOPE:

This SOP covers the procedure for Chemical Analysis of water

- RESPONSIBILITY/ACCOUNTABILITY:

- QC analyst shall be responsible for testing water

- QC Head or designee approves the analytical results and releases water for production use.

4.0 ABBREVIATIONS AND DEFINITIONS

PPE– Personal Protective Equipment

- PROCEDURE

APPEARANCE:

- Take 30ml of the water sample in a clean and clear test tube.

- View down the vertical axis in a bright light against black and white background.

- Expected observation: Clear and colourless.

CONDUCTIVITY:

- Take 50ml of water sample in a beaker.

- Immerse the conductivity meter probe.

- Invigorate and read.

- Expected observation: Not more than 10s/cm at room temperature.

ACIDITY / ALKALINITY:

- Pipette 20ml of water sample into a clean test tube.

- Add 0.05ml of Phenol red solution.

- Expected observation: Not more than 0.1ml/0.15ml of 0.01M Sodium hydroxide / 0.01M Hydrochloric acid is required to turn the colour to red/yellow.

OXIDISABLE SUBSTANCES:

- Take 100ml of the water sample in a 250ml beaker.

- Add 10ml of dilute sulphuric acid and boil.

- Add 0.2ml of 0.02M Potassium permanganate.

- Boil further for 5 minutes.

- Expected observation: The solution remains faintly pink.

RESIDUE ON EVAPORATION:

- Take 100ml of the water into a crucible.

- Evaporate to dryness on water bath.

- Dry in oven at 100 – 105oC.

- Expected observation: Not more than 3mg.

LIMIT TEST:

Heavy Metals

- Heat 200ml of sample in a glass evaporating dish on a water bath until the volume in reduced to 20ml.

- Take 12ml of the concentrated sample in a test tube and mixture of 10ml of lead standard solution (1ppm) with 2ml of concentrated sample in another test tube.

- Add 2ml of acetate buffer 3.5 to each test tube.

- Mix and add 1.2ml of Thioacetamidle reagent to each test tube.

- Mix immediately and allow them to stand for 2 minutes.

- Expected observation: The standard solution is more intensely coloured than the test solution.

Ammonium

- Place 20ml of the water sample in a clean test tube and a mixture of 4ml of ammonium standard solution (1ppm) with 16ml of ammonium free water in another test tube.

- Add 1ml of alkaline potassium tetraiodomercurate solution to each test tube.

- Examine the solutions down the vertical axis of the tubes.

- Expected observation: The test solution is not more intensely coloured than the standard solution.

Nitrates

- Place 5ml of the water sample in a test tube and a mixture of 4.5ml nitrate free water with 0.5ml of nitrate standard solution (2ppm) in another test tube.

- Add 0.1ml of diphenylamine solution to each.

- Add 5ml of nitrogen free sulphuric acid drop wisely with shaking.

- Transfer the tubes to water bath at 50oC.

- Expected observation: The solution shows no change in appearance for atleast 15 minutes. Any subsequent colour is not more intense than standard solution.

Sulphates

- Place 10ml of the water sample in a test tube.

- Add 0.1ml of dilute hydrochloric acid.

- Add 0.1ml of barium chloride solution.

- Expected observation: The solution shows no change in appearance for at least 1 hour.

Chlorides

- Place 15ml of the water sample in a test tube and mixture of 1.5ml of chloride standard solution (5ppm cl-) with 13.5ml of water in another test tube.

- Add 1ml of dilute nitric acid to each test tube.

- Add 0.2ml of silver nitrate solution to each test tube.

- Expected observation: The solution shows no change in appearance for at least 15 minutes and any subsequent colouration is not more intense than standard solutions.

Calcium and Magnesium

- Place 50ml of sample in a beaker.

- Add 1ml of ammonium buffer solution 10.0.

- Add 50mg of mordant black II triturate.

- Add 0.5ml of 0.01M sodium edetate.

- Expected observation: A pure blue colour is produced.

Materials, Equipment, Apparatus and Reagents

- PPE

- Test tube

- Beaker

- pipette

- Conical flask

- 0.01M Sodium hydroxide

- 0.01M Hydrochloric acid

- Sulphuric Acid

- Barium chloride

- Ammonium buffer

- Phenol red

- Nitric Acid

- Silver Nitrate

- 0.01M Sodium Edetate

- Potassium tetraiodomercurate

- Thioacetamidle

- Diphenylamine solution

- Crucible

- Oven

- 0.02M Potassium permanganate

Safety, Health and Environment Precautions

- Be prompt to clean any chemical spillage

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

In House

- TRAINING INVOLVEMENT

Quality Manager

Quality control staff

All sectional heads in the quality department

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |