- OBJECTIVE:

This SOP describes the process of testing for fertility of microbiological media

- SCOPE:

This SOP covers all microbiological media used in the microbiological laboratory.

- RESPONSIBILITY/ACCOUNTABILITY:

- Microbiologist/ Assistant shall be responsible for conducting the growth promotion test on all media.

- The Quality Control Head ensures the strict adherence to this procedure.

- Abbreviations and Definitions

- GPT: Growth Promotion Test

- NLT: Not Less Than

- NMT: Not More Than

- PROCEDURE

General:

- It should be observed/assumed that samples and microorganisms with which you work in the microbiological laboratory can cause illness and disease. Great care should therefore be taken in handling cultures, slides and all materials that may have come in contact with living organisms.

- All live cultures shall be segregated from areas used for sample testing

- Positive control dilutions and inoculation shall be performed in a biological safety cabinet

- Surfaces, work areas shall be sanitized immediately after use.

- Gloves shall be used when working with live cultures.

- Growth promotion test shall be performed whenever a new medium is received, Batch No. is changed, and Make is changed

- After procuring the media, check the following details i.e. manufacturing date, Expiry date and Batch No.

- Store all the media in a cool dark place at NMT 25 ± 5°c

- Before taking the media for use, perform the growth promotion test using the organisms as mentioned in guide Acceptance Criteria for the Growth promotion Test for Culture Media

BML-QDG-MC-002-G01

- Whenever there is an approved and ‘in use’ medium available, inoculate it with the same organism

- If the growth promotion qualities of the media are not same as compared to the previously approved batches, then discard that media.

- Growth promotional test is carried out at the syrup micro laboratory after appropriate line clearance done.

For newly received agar media

- Prepare media to be tested as per sop for media preparation (BML-QDG-MC-001)

- from test tube dilution having cell count of 10 to 100 cells, transfer 1.0ml of the aliquot into two sterile petri dishes and pour the media of New container (B)

- swirl the plates and allow to solidify

- similarly, transfer 1.0ml of the culture of having a cell count of 10 to 100 cells in dilution into two sterile petri dishes and pour the prepared previous container media (A)

- swirl the plates and allow to solidify.

- Incubate the plates at respective temperatures, after incubation count the colonies and compare the count with the previous container results.

- After incubation record the results into the Comparative test report form BML-QDG-MC-002-F01

- The recovery in the growth promotion test for new container must be within factor 2 of the actual inoculum concentration obtained for the previous container e.g. if the standard culture inoculum used has count 100cfu/ml then the recovery should be within 100/2 = 50 cfu/ml and 100 x 2 =200 cfu/ml.

For agar media:

- Prepare inoculum of known concentration (approximately 10 to 100 cells/ml) of the organism and then transfer 1ml of the inoculum to the plate followed by pouring with the agar media

- After incubation record the results into the Growth promotion test for Agar media BML-QDG-MC-002-F02

For selective agar media:

- Carryout growth promotional test by streaking specified organisms on the plate, observe for the characteristics of colonies, and record the result into the Growth promotion test for Agar media BML-QDG-MC-002-F02

In case of liquid media or broth:

- Inoculate 1ml of the inoculum in 10ml of media by using the organisms as mentioned in the guide of Acceptance Criteria for the Growth promotion Test for Culture Media BML-QDG-MC-002-G01

- Incubate the media respectively

- Observe the growth and record the result in Growth promotion test for Liquid media

BML-QDG-MC-002-F03

- Liquid media are suitable if should show clearly visible growth in media

- Put an uninoculated bottle/ plates of media as a negative control to confirm the sterility of the media

- For the growth promotion test of daily-prepared media, if more organisms are prescribed for the test, in that case, minimum of two bacteria and one mold/yeast shall be used for tests on daily rotation basis.

- REFERENCE/ENCLOSURE

- Guide for the Growth promotion Test for Culture Media - BML-QDG-MC-002-G01

- Newly received agar media Comparative test report form - BML-QDG-MC-002-F01

- Growth promotion test for Agar media record form - BML-QDG-MC-002-F02

- Growth promotion test for Liquid media record form - BML-QDG-MC-002-F03

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORIES |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |