- OBJECTIVE:

This procedure is to lay down a procedure for proper disposal of used media

- SCOPE:

This procedure is applicable for used media disposal

- RESPONSIBILITY/ACCOUNTABILITY:

- Microbiologists are responsible for the operation and disposal of the culture media.

- The Quality control head ensures the procedures outlined are followed.

4.0 ABBREVIATIONS AND DEFINITIONS

- PPE: Personal Protective Equipment

- PROCEDURE

FOR MICROBIAL MEDIA PLATES

- Wear rubber hand gloves and nose mask

- Arrange the closed petri dishes that contain the culture media in to canister

- Place the canister into bucket  autoclave

- Autoclave at 1210c, 15psi for 30 minutes.

- After autoclaving, add 5% of sodium hypochlorite solution to the media and dilute, drain into the sink under running tap water.

- Keep the empty glass Petri-dishes for 1hr in 5% sodium hypochlorite solution, then wash the Petri-dish with soap solution and rinse thoroughly with clean water

- Perform final rinsing of the petri dishes using 70% Isopropyl Alcohol  with Water for Injection(WFI)

- Stack the dish base and lids together

- Wrap each stack in aluminum foil

- Seal lightly to prevent contamination but allow stream penetration

- Place wrapped dishes in the autoclave

- Autoclave at 1210c, 15psi for 15 minutes

- Dry them in oven(NLT 800c)for re-use

- Disposable plates should be discarded after keeping them for 1hr in 5% sodium hypochlorite.

FOR CULTURED TUBES

- Autoclave the microbial media tubes at 15ps, 1210c for 30 minutes.

- Drain the contents of tubes into the sink under running tap water.

- Keep the empty tubes for 1 hrs in 5% sodium hypochlorite solution, then wash the Petri-dish with soap solution and rinse thoroughly with clean water.

- Perform final rinsing of the tubes using 70% Isopropyl Alcohol with Water for Injection (WFI)

- Wrap each tubes in aluminum foil

- Place wrapped tubes in the autoclave

- Autoclave at 1210c, 15psi for 15 minutes

- Dry them in oven(NLT 800c)for re- use

MATERIALS AND EQUIPMENT:

- PPE

- Cultured media Petri-dishes

- Cultured tubes

- 5% sodium hypochlorite solution

- Liquid soap solution

- WFI water

- 70% IPA

- Oven

Safety, Health and Environment Precautions

Wear appropriate PPE

- REFERENCE/ENCLOSURE

USP

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORIES |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |