| S.N | AUDIT FINDINGS (OBSERVATION) | ROOT CAUSE ANALYSIS | CORRECTIVE ACTION | INDICATORS FOR COMPLETION | TIMELINE | RESPONSIBILITY | CAPA STATUS |
| --- | --- | --- | --- | --- | --- | --- | --- |
| GENERAL Q.C |  |  |  |  |  |  |  |
| 1 | Quality control testing methods was not available as at the time of inspection |  |  |  |  |  |  |
| 2 | A few calibrated standard weights were available for sighting, as other higher weight was not sighted. |  |  |  |  |  |  |
| 3 | A few calibrated standard weights were available for sighting, as other higher weight was not sighted. |  |  |  |  |  |  |
| 4 | Equipment cleaning record were not available |  |  |  |  |  |  |
| 5 | Though the laboratory was equipped with necessary equipment to carry out daily activities, the staffs believe Atomic absorption spectroscopy for determination of chemical elements is desirable. (i.e. to analyse elements in solution) |  |  |  |  |  |  |
| 6 | Test for disinfectant effectiveness were found not to be updated. |  |  |  |  |  |  |
| CLEANING & GENERAL HOUSE KEEPING |  |  |  |  |  |  |  |
| 1 | SOP for cleaning does not state time of cleaning |  |  |  |  |  |  |
| 2 | Pest control mapping was not found |  |  |  |  |  |  |
| 3 | Top of work benches were found with liquids (may be water or chemical) indicating poor cleaning of benches. |  |  |  |  |  |  |
| 4 | The department cabinets were not adequately cleaned while broken and unbroken glassware were found in some cabinets. |  |  |  |  |  |  |
| 5 | Records were found scattered in shelves, without proper arrangement |  |  |  |  |  |  |
| RAW MATERIAL TESTING |  |  |  |  |  |  |  |
| 1 | Specification and method of analyses were not found, while staff claims they are under review. |  |  |  |  |  |  |
| 2 | Raw materials register and notebook though sighted were not paginated, while most activities were not signed by supervisor. |  |  |  |  |  |  |
| 3 | Working standards not seen. |  |  |  |  |  |  |
| IN-PROCESS & FINISHED PRODUCT TESTING |  |  |  |  |  |  |  |
| 1 | Specification & testing methods for in-process and test methods not validated nor sighted |  |  |  |  |  |  |
| 2 | Test analysis notebook sighted but not robust as notebook was not signed |  |  |  |  |  |  |
| 3 | In-house calibration records not presented. Bottle calibration equipment record deficient |  |  |  |  |  |  |
| 4 | Q.C. Instrument not fully labelled, refer to that plate equipment not labelled. |  |  |  |  |  |  |
| 5 | Calibration tags missing in some equipment e.g. weigh scale, digital Polari meter and conductivity meter. |  |  |  |  |  |  |
| 6 | No record as per laboratory weight supplied. |  |  |  |  |  |  |
| MICROBIOLOGY LAB |  |  |  |  |  |  |  |
| 1 | Although did not have access to Micro lab, all questions were answered responded orally. - Good laboratory practice not fully adhered to, as supervisor did not sign notebook |  |  |  |  |  |  |
| 2 | Due to inaccessibility, area temp & monitory record could not be checked |  |  |  |  |  |  |
| 3 | Analytical methods under review. Analyst notebook available but supervisor did not sign all papers. |  |  |  |  |  |  |
| 4 | Set environmental monitoring standard & program was not sighted |  |  |  |  |  |  |
| 5 | Effectiveness of Lamel flow was not checked |  |  |  |  |  |  |
| 6 | Maintaining record was not provided. |  |  |  |  |  |  |
| STABILITY LAB |  |  |  |  |  |  |  |
| 1 | Accelerated stability record was not updated. Stability chamber back up not working as the time of assessment. |  |  |  |  |  |  |
| GENERAL |  |  |  |  |  |  |  |
| 1 | SOP could not be seen , as the team was referred to the documentation office for GLP, SOP should be placed near the analyst work beech |  |  |  |  |  |  |
| 2 | Clearing record of each equipment was not available |  |  |  |  |  |  |
| 3 | Cleaning record of each equipment was not available |  |  |  |  |  |  |
| 4 | Extractor piping found to be dirty & not clean |  |  |  |  |  |  |
| 5 | Power backup of stability chamber not working |  |  |  |  |  |  |
| 6 | Cupboards were found not tidy |  |  |  |  |  |  |
| 7 | Laboratory signature atlas not found |  |  |  |  |  |  |
| 8 | Danger signature (Flammable, Hazard) not sighted |  |  |  |  |  |  |
| 9 | Most stools in lab were found to be wooden. |  |  |  |  |  |  |
| 10 | Gas mask / respirator not available for each staff |  |  |  |  |  |  |
| 11 | Gowning procedure in QC not robust |  |  |  |  |  |  |
| 12 | Instrumentation lab epoxy floor not coving to the wall. |  |  |  |  |  |  |
| 13 | Reference to training done (on 03/08/2024), some attendant did not sign training sheet. |  |  |  |  |  |  |