- OBJECTIVE:

This SOP describes the process for the non-destructive integrity testing of sterilizing-grade filters to confirm their structural integrity and ability to retain contaminants filters integrity using the bubble point test method

- SCOPE:

This SOP is applicable to all sterilizing grade filters (e.g., 0.22 µm, 0.45 µm) used in the production of intravenous fluid in BIOMEDICAL Ltd.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Microbiologist: Responsible for performing the filter integrity test as per the SOP and recording the results.

- In-process Quality Assurance (IPQA): Responsible for verifying the test results, ensuring compliance, and final batch disposition.

- Engineering/Maintenance: Responsible for the calibration and maintenance of the integrity testing equipment.

- ABBREVIATION AND DEFINITIONS

Nil

- PROCEDURES

5.1 General

- To be carried out on every fresh supply of filters received by the stores and sampled by the quality assurance department

- To periodically validate membrane filters integrity during operations

- To be Quality Control department shall arrange for sampling of materials as soon as notification note is received from concern department.

- Sampling to be done by a trained competent person from the quality control department.

- Testing is typically performed both before (pre-use) and after (post-use) the filtration process. Post-use testing is a regulatory requirement for batch release.

  - Preparation

- Check the filter manufacturer's data sheet for critical parameters, including pore size, membrane material, recommended wetting fluid, and the expected pass/fail values (e.g., minimum bubble point pressure or maximum diffusion flow rate).

- Ensure the integrity test machine or manual test apparatus (pressure source, calibrated pressure gauge, etc.) is clean and calibrated.

- Install the filter cartridge into its clean housing, ensuring all O-rings and seals are properly seated to prevent leaks.

  - Wetting (for Bubble Point)

- A hydrophilic filter used is wet with sterile Water for Injection (WFI). Slowly fill the housing with WFI, ensuring all air is bled from the system through the vent valve. A static high-pressure hold may be used to ensure complete wetting.

5.4 Test Method

- Fix the filter sample to be tested in the filter holding unit

- Open the water line to inlet of the filter holding unit and wet the filter by passing about 20 liters of WFI water into the holding unit.

- Connect air from a pressure air- line into the inlet of the filter holding unit.

- Connect the out let of the filter holding unit to rubber tubing.

- Immersed the outlet rubber tubing in water.

- Slowly raise the air speed by opening the pressure air out let valve.

- Note the pressure in the pressure gauge when there is continuous stream of bubbles of air in the bucket that the outlet rubber tubing is immersed

- If the pressure is 26psi and above, the membrane filter passes the test and can be used for filtration purpose

- If the pressure is below 26 psi, then discard the membrane filter and fix a new one. Repeat the same procedure from step 2

- If on second trial of the sample filter with another sample it fails the test, repeat the test with another sample from another pack

- If sample from another pack fails the test, the entire lot stands rejected. However if it pass, then repeat the test once again to confirm

- Record your findings and observations in the Filter integrity test record book

- The recorded bubble point pressure must be at or above the manufacturer's specified minimum value

- Record all test results, including filter lot number, operator name, date, time, and test conditions (e.g., pressure, flow rate, temperature) in the relevant record sheets.

Note: If a filter fails the integrity test, follow the established troubleshooting procedures, which may involve re-wetting and retesting, but ultimately may require replacing the filter and investigating the cause of failure.

- Required materials:

- Plastic bucket

- Rubber Hose

- Filter Cartridges

- Safety, Health and Environmental Precautions:

- Always wear PPE

- Care should be taking when handling pressurize air hose

- Do not slip water on the floor, if any water please mop up immediately

- REFERENCE AND ENCLOSURE

Nil

7.0 TRAINING INVOLVEMENT

| IPQA Officers | Microbiologist |
| --- | --- |
| QA Doc. Unit staff | QA/Engineering staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |