- OBJECTIVE:

This SOP describes the procedure to confirm the autoclave's ability to achieve and maintain adequate steam sterilization conditions (temperature, time, and steam) for a typical load using chemical indicator strips.

- SCOPE:

To provide written guidelines for establishing the efficacy of the Autoclave using chemical strips in the Microbiological Laboratory in Infusion/Syrup in Biomedical Ltd.

- RESPONSIBILITIES/ACCOUNTABILITIES:

Quality Assurance Officer, Chemist, Microbiologist, Quality Assurance Manager, Superintendent Pharmacist/ Factory Manager

- ABBREVIATION AND DEFINITIONS

Nil

- PROCEDURES

- Take three chemical trips to a bag of infusion meant to be sterilized using paper tape.

- Place one strip on the top tray, one also at the middle and other to the lower tray.

- Allow the product to be autoclaved at 110oC, 1.7 bars and 60 minutes.  After the cycle, remove the strip sand check.

Observation

- The colour will change from white to black and from unsafe side to safe side.

Conclusion:

Autoclaving is validated when the process can suitably render a constant result

Materials

- Autoclave tape

- Autoclave

- REFERENCE AND ENCLOSURE

- In-house Method

7.0 TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff | All staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |