- OBJECTIVE:

This SOP describes the procedure of operation, Calibration and maintenance of Analytical balance in the Quality Control Department.

- SCOPE:

To provide written procedure for calibration of Analytical weighing balance in the Quality Control Department.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- QC Analysts shall be responsible for the Operation, calibration and maintenance of the Instrument/Equipment as per the SOP.

- Maintaining the history card, instrument usage logbook as per the SOP.

- QC Head/ Manager shall be responsible to ensure that the instrument is operated , calibrated and maintained and properly documented as per the SOP.

- QA shall be responsible to ensure the implementation of the system as per the SOP.

- QAM shall be responsible for the approval of the SOP.

- ABBREVIATION AND DEFINITIONS

- Cal Int – internal calibration

- Cal done – Calibration Done

- IPA – Isopropyl Alcohol

- NA- Not Applicable

- NMT – not more than

- NIST – National Institute of Science and Technology.

- PROCEDURES

GENERAL NOTES NAD PRECAUTIONS:

- The balance door should not be open during weighing.

- Do not expose hygroscopic material more time in the air during weighing.

- Air currents should not present near the balance.

- The balance must be properly leveled.

- Do not cause vibration during laboratory operations.

- Pick up the weight with a forceps; place it carefully on the pan.

- Place the weight at the center of the balance pan to eliminate the corner-weighing difference.

- Do not overload the balance pan more than the capacity of the balance

- Many chemicals such as salts (corrosive) and materials of this nature should not spill on the balance pan or inside the balance housing.

- Before weighing or calibration, the level indicator which should be in the center.

- The balance should be calibrated at least once daily or preferably before every weighing operation.

- First clean the balance with a soft cloth.

- Ensure that the scale is at zero, if not adjust the dial until the arrow on the scale rests at ZERO.

- Put the standard weights starting from the lowest weight available to the highest weight.

- Record the weight displayed

- The reading should not vary more than+ 0.1% of the weight being measured x least count on the balance.

- If the readings are satisfactory according to the above use the balance for weighing operation otherwise handover the balance to the Maintenance department for repair.

- The weight used for calibration shall be calibrated against standard weight traceable to NIST.

- Record all data in the calibration sheet, ensuring that required persons sign against the calibration performed.

- REFERENCE AND ENCLOSURE

ENCLOSURE 1- Operation Manual supplied by the Manufacturer.

ENCLOSURE 2- Weighing Balance Calibration Log - BML-QDG-QA-032-F01

USP<41> weights and balances

SOP for maintenance of Laboratory instruments

USP <1251> weighing on an Analytical balance.

7.0 TRAINING INVOLVEMENT

| Quality Control Manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | QC Staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |

ENCLOSURE 1

Weighing Balance Calibration Log

|  | Biomedical Limited, No. 1 Ohimege Road, Industrial Estate, Ilorin, Kwara State, PMB 1449 | Doc.No.: BML/QDF/QA/032 |
| --- | --- | --- |
|  | QUALITY ASSURANCE DEPARTMENT EQUIPMENT CALIBRATION LOG ANALYTICAL WEIGHING BALANCE | Version No.: 02 |

Doc. No.; BML/QDG/QA/032/F01

| Date | Standard weight used (g)(kg) | Readings | Deviation | Specification (+ 0.1%) | Remark | Done By | Checked by |
| --- | --- | --- | --- | --- | --- | --- | --- |

_____________________

Q.A Manager

Sign /Date