- OBJECTIVE:

This SOP describes the procedure to be followed in Calibration of Incubator

- SCOPE:

To provide written procedure for calibration of Incubator

- RESPONSIBILITIES/ACCOUNTABILITIES:

QA Officer, Microbiologist, QA Manager

- ABBREVIATION AND DEFINITIONS

QC – Quality Control

- PROCEDURES

CALIBRATION PROCEDURE:

- Check the thermostat if it actually set to the required temperatures (i.e. 250C for fungi growth and 370C for Bacteria).

- Check if the metal shelves are cleaned.

- Make sure that the door of the incubator is properly otherwise temperature will leak and accuracy will not be maintained.

- Place 5 calibrated thermometer inside the incubator in different places as per the location diagram attached as Enclosure 2

- Operate the incubator as per the standard operating procedure –BML/MCB/

- Document your readings from each calibrated thermometer at an interval of 10minutes and record the results up to 24hours

- Report to QC Manager immediately if there is lapses

- Frequency – once in 3 month

- REFERENCE AND ENCLOSURE

- ENCLOSURE 1- Calibration Record of Incubator- BML-QDG-QA-048-FO1

- ENCLOSURE II: LOCATION DIAGRAM OF EACH THERMOMETER

7.0 TRAINING INVOLVEMENT

| Quality Control Manager | Quality Control Microbiologist |
| --- | --- |
| QA Doc. Unit staff | Lab Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |

Enclosure 1

Calibration record of Incubator

Sop Ref.No. BML/QDG/QA/048

Doc .No: BML/ QDG/QA/048/F01

Frequency: once in 3 month

| Date | Set temperature (oc) | Thermometer no. | Thermometer Temp. readings | Incubator temp. Reading | Deviation ( +2) |
| --- | --- | --- | --- | --- | --- |

Done by:

Checked by: