- OBJECTIVE:

The purpose of this SOP is to describe the procedure for the calibration of digital pH meter.

- SCOPE:

To provide written procedure for calibration of Digital pH Meter

3.0 RESPONSIBILITY/ACCOUNTABILITY:

Water Treatment Officer, Chief Engineer, Chemist, Q.A Manager, Superintendent Pharmacist/ Factory Manager.

- ABBREVIATIONS AND DEFINITIONS:

N/A

- PROCEDURE:

CALIBRATION:

- Ensure that the electrode of the pH meter is connected properly to the display unit and the unit is plugged to the power supply.

- Switch on the power supply to the unit.

- Rinses the electrode with distils water and wipe it with tissue paper gently.

- Dip the combined electrode in a 7 pH buffer solution.

- Set the “TEMPERATURE C” control to the temperature of the selected buffer solution.

- Set the function selector switch to pH position and adjust “CALIBRATE “control until the meter displays the precise pH value of the buffer solution.

- Now set the function selector switch to STAND BY.

- Remove the electrode from the buffer solution. Wash with distilled or deionized water.

- Dip the combined electrode into another buffer solution (Set 4 pH Value).

- Set the “TEMPERATURE C” control to the buffer solution temperature.

- Set the function selector switch to pH position. Adjust the “SLOPE” correction control, until the meter displays the pH value of the selected buffer solution. Check that the correct readings are obtained with both solutions without further adjustment.

- If the readings are not in the stipulated limits, repeat the calibration again.

- If still no good results, inform the QC Manager about it. Do not use the pH meter when the calibration is not good.

- If the result of calibration is ok, then go for measuring the pH of the sample.

- The pH meter should be calibrated at least once a day

REFERENCE/ENCLOSURE:

In-house

- TRAINING INVOLVEMENT

| Chief Engineer | Engineer-Incharge, Operators/Assistants. |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | WATER TREATMENT UNIT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |