- OBJECTIVE:

To provide the instructions for cleaning and operation of the Pharma Lab autoclave.

- SCOPE:

To provide written guidelines for cleaning and operation of Pharma Lab Autoclave.

- RESPONSIBILITIES/ACCOUNTABILITIES:

Production Pharmacist, Production Supervisors and Operators.

- ABBREVIATION AND DEFINITIONS

BMR- Batch Manufacturing Record.

- PROCEDURES

Cleaning Procedure:

- Ensure that all cleaning equipment are clean i.e. buckets and soft cloth.

- Make sure all reject pouches in the autoclave are removed.

- Fill all buckets needed with water sufficient quantities of liquid soap to it.

- Dip a clean napkin into the solution and clean the autoclave’s interior and exterior thoroughly.

- Then use the clean fresh water to rinse the soap solution from the autoclave and make sure the solution is properly drained.

- Make sure all the exterior part of the autoclave is dried with the napkin provided then wash and dry all the cleaning equipment.

- Document the cleaning process.

NOTE:

- Do not use detergent in cleaning the autoclave.

Operation Procedure:

- Ensure that the boiler and air compressor are on, working properly and their valves open.

- Remove the ‘unsterilized’ tag from the trolleys.

- Wheel the loaded trolley into the autoclave and close the door.

- Switch on the main starter, Auto manual, Steam and Condensate button to start sterilization.

- Monitor the temperature and pressure readings and record it in the autoclave log sheet book.

- When the temperature reaches 116°C; pressure, 1.5 Bar; allow to sterilize for 40 minutes; then allow the air and water cooling to complete the cooling phase.

- Once the sterilization cycle is complete, open the doors and wheel out the sterilized product from the autoclave.

- Collect the autoclave reading print out and attach to the BMR.

- Set aside quality control samples and move the sterilized product to the offloading section, sterilized product area.

Materials, Equipment, Apparatus and Reagents

- WFI water.

- Soft cloths.

- Mops.

- Buckets.

- Soap.

- REFERENCE AND ENCLOSURE

User Manual, In House Method.

7.0 TRAINING INVOLVEMENT

| Production manager | Production pharmacist |
| --- | --- |
| Sectional heads | All production staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | PRODUCTION DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |