- OBJECTIVE:

To provide the instructions for cleaning and operation of the Pharma Lab autoclave

- SCOPE:

To provide written guidelines for cleaning and operation of Pharma Lab Autoclave

- RESPONSIBILITIES/ACCOUNTABILITIES:

Production Pharmacist, Production Supervisors and Operators

- ABBREVIATION AND DEFINITIONS

BMR- Batch Manufacturing Record

- PROCEDURES

Cleaning Procedure:

- Ensure that all cleaning equipment are clean i.e. buckets and soft cloth.

- Make sure all reject pouches in the autoclave are removed.

- Fill all buckets needed with water sufficient quantities of liquid soap to it.

- Dip a clean napkin into the solution and clean the autoclave’s interior and exterior thoroughly.

- Then use the clean fresh water to rinse the soap solution from the autoclave and make sure the solution is properly drained.

- Make sure all the exterior part of the autoclave is dried with the napkin provided then wash and dry all the cleaning equipment.

- Document the cleaning process

NOTE:

- Do not use detergent in cleaning the autoclave.

Operation Procedure:

- Ensure that the boiler and air compressor are on, working properly and their valves open.

- Remove the ‘unsterilized’ tag from the trolleys

- Wheel the loaded trolley into the autoclave and close the door.

- Switch on the main starter, Auto manual, Steam and Condensate button to start sterilization.

- Monitor the temperature and pressure readings and record it in the autoclave log sheet book.

- When the temperature reaches 116°C; pressure, 1.5 Bar; allow to sterilize for 40 minutes; then allow the air and water cooling to complete the cooling phase.

- Once the sterilization cycle is complete, open the doors and wheel out the sterilized product from the autoclave

- Collect the autoclave reading print out and attach to the BMR

- Set aside quality control samples and move the sterilized product to the offloading section, sterilized product area.

Materials, Equipment, Apparatus and Reagents

- WFI water

- Soft cloths

- Mops

- Buckets

- Soap

- REFERENCE AND ENCLOSURE

User Manual, In House Method

7.0 TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff | All staff |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |