- Objective:

To provide written guidelines for cleaning and Operation of analytical balance.

- Scope:

This SOP covers the process of cleaning and Operation of analytical balance used in

the chemical laboratory.

3.0 Responsibility/accountability:

- Laboratory Assistant, Quality Control Chemist, Quality Assurance Officer,

Quality Assurance Manager, Superintendent Pharmacist/Factory Manager.

4.0 Abbreviations and Definitions

cGMP – Current Good Manufacturing Practices

SOP – Standard Operating Procedure

PPE- Personal Protective Equipment

UPS- Uninterruptible Power Supply

IPA- Isopropyl Alcohol

- Procedure

#### 5.1 Safety Precautions

- Always switch OFF the analytical balance and main power supply before cleaning.

- Do not move or reposition the balance while cleaning.

- Wear appropriate PPE (lab coat, gloves, safety goggles).

- Handle calibration weights and internal components with care.

#### 5.2 Cleaning Procedure

- Open the side or top glass door of the balance.

- Clean the weighing pan and cubicle interior using a dry, lint-free cloth.

- Wipe the top and side glass with a clean dry cloth.

- Re-clean the pan and weighing chamber using a cloth dampened with 70% IPA.

- Remove any localized stains.

- Dry all cleaned surfaces thoroughly.

- Calibration weights should be cleaned before and after each use with a dry lint-free cloth.

- Record the activity in the Analytical Weighing Balance Usage Logbook (BML-QDF-CL-004-F01)

#### 5.3 Operation Procedure

- Switch ON the main power and the UPS connected to the balance.

- Power the balance ON by pressing and holding the 0/T button.

- Allow the unit to warm up for at least 15 minutes before use.

- Perform calibration as per Calibration SOP (BML-QDF-QA-023).

- Ensure maximum load does not exceed 150g.

- Always use a weighing dish and spatula when transferring materials.

- Use the 0/T (tare) button to reset the balance to zero before weighing.

- After use:

  - Ensure the balance is clean and free of spills.

  - Switch OFF the power and disconnect from the main power supply.

  - Close the glass panels.

- Record the activity in the Analytical Weighing Balance Usage Logbook (BML-QDF-CL-004-F02)

6.0 REFERENCE/ENCLOSURE

IN-HOUSE

ENLCOSURE 1- ANALYTICAL BALANCE CLEANING LOGBOOK- BML-QDF-CL-004-F01

ENCLOSURE 2- ANALYTICAL BALANCE USAGE LOGBOOK- BML-QDF-CL-004-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QC Analyst | Laboratory Assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |