- Objective:

To describe the process of Cleaning and operation of Handheld pH Meter

- Scope:

This SOP covers the Cleaning and operation of the handheld pH Meter

3.0 Responsibility/accountability:

Quality Control Chemist, Quality Control Microbiologist, Quality Assurance Manager and the Superintendent Pharmacist/ Factory Manager

4.0 Abbreviations and Definitions

PPE - Personnel Protective Equipment

- Procedure

### Cleaning Procedure

Perform cleaning weekly or after use in oily samples:

- Rinse the electrode in distilled water.

- Soak the electrode in cleaning solution (or equivalent) for 15–30 minutes.

- Rinse thoroughly with distilled water after cleaning.

- Blot gently with a soft tissue.

- Store in electrode storage solution (not distilled water).

- Record the activities done in the Handheld pH meter Cleaning Logbook (BML-QDF-CL-008-F01)

### Operation procedure

### Meter Preparation

- Remove the protective cap from the electrode.

- Rinse the electrode with distilled water and gently blot dry.

- Inspect the probe for any damage or air bubbles. Do not rub the glass bulb.

- Turn on the meter using the power button.

Calibration (to be done daily or before use)

- Prepare standard buffer solutions (e.g., pH 4.01 and 7.01).

- Rinse the electrode with distilled water.

- Immerse the electrode in pH 7.01 buffer. Wait for the reading to stabilize.

- Press and hold the “CAL” or “MODE” button (depending on the model) to initiate calibration.

- Once confirmed, rinse the electrode and immerse in the pH 4.01 buffer.

- Allow reading to stabilize and confirm second point of calibration.

- If available, perform a 3-point calibration using pH 10.01.

- Exit calibration mode when prompted.

### Sample Measurement

- Rinse the electrode with distilled water.

- Immerse the electrode in the sample solution.

- Gently stir the sample and wait for the reading to stabilize.

- Record the pH value.

- Rinse the electrode between samples.

- Record the activities done in the Handheld pH meter Usage Logbook (BML-QDF-CL-008-F02)

Required Materials:

-	Personnel Protective Equipment’s (PPE)

- Distilled water

- Reagent bottles

6.0 REFERENCE/ENCLOSURE

IN-HOUSE

ENCLOSURE 1- HANDHELD pH METER CLEANING LOGBOOK – BML-QDF-CL-008-F01

ENCLOSURE 2- HANDHELD pH METER USAGE LOGBOOK – BML-QDF-CL-008-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QC Analyst | Laboratory Assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |