- Objective:

To provide written guidelines for cleaning and operation of UV Spectrophotometer.

- Scope:

This SOP covers the process of cleaning and operation of UV Spectrophotometer.

3.0 Responsibility/accountability:

Quality Control Chemist, Quality Assurance Manager, Superintendent Pharmacist/ Factory Manager.

4.0 Abbreviations and Definitions

cGMP – Current Good Manufacturing Practices

SOP – Standard Operating Procedure

PPE- Personal Protective Equipment

- Procedure

#### 5.1 Cleaning Procedure:

- Ensure the instrument is switched off and disconnected from the power supply before cleaning.

- Routine cleaning should be carried out to prevent dust accumulation.

- Once weekly, carry out a thorough cleaning procedure as follows:

  - Open the sample chamber and clean the cuvette holder with a soft lint-free cloth.

  - Clean the outer surfaces of the equipment with a dry lint-free cloth.

  - Carefully lift the instrument (if necessary) and clean the base area before returning it to its original position.

  - Avoid use of any corrosive or abrasive cleaners; use 70% IPA or ethanol if necessary.

  - Allow the equipment to dry before reconnecting to power.

- Record the activities in the UV-spectrophotometer Cleaning Logbook (BML-QDF-CL-011-F01).

#### 5.2 Operation Procedure:

##### Initial Set-up:

- Ensure the rear voltage selector switch is correctly set for the local power supply.

- Plug in and switch on the spectrophotometer.

- Allow 30 minutes for system stabilization and lamp warm-up.

##### Measurement Procedures:

A. Blank Reading:

- Fill a clean cuvette with blank solution (usually the solvent used for the sample).

- Wipe the cuvette clean and insert it into the cell holder.

- Close the sample chamber lid.

- Select the required wavelength.

B. Absorbance (ABS) Measurement:

- Press the T/A/C/F keypad to select Abs mode.

- Press 100%T/0A to zero the instrument.

- Remove the blank, insert sample cuvette, and read the absorbance directly.

C. Percent Transmittance (%T) Measurement:

- Select %T mode using the T/A/C/F keypad.

- Press 100%T/0A to standardize at 100%.

- Insert the sample cuvette and read the % transmittance directly.

D. Concentration Measurement:

- Calibrate the instrument as described above.

- Remove blank and insert a known standard solution in a cuvette.

- Record the absorbance or transmittance.

- Repeat measurements with sample solution(s).

- Record the activities in the UV-spectrophotometer Usage Logbook (BML-QDF-CL-011-F02)

Post-operation Steps:

- Remove and clean all cuvettes.

- Turn off the instrument.

- Disconnect from the power supply if not in use for an extended period.

Materials, Equipment, Apparatus and Reagents

- PPE, soft lint cloth.

6.0 REFERENCE/ENCLOSURE

In House

ENCLOSURE 1- UV-SPECTROPHOTOMETER CLEANING LOGBOOK- BML-QDF-CL-011-F01

ENCLOSURE 2- UV-SPECTROPHOTOMETER USAGE LOGBOOK- BML-QDF-CL-011-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| QC Analyst | Laboratory Assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |