- Objective:

To provide written guidelines for cleaning and Operation of Flame photometer

- Scope:

This SOP covers the process of cleaning and Operation of Flame photometer used in the chemical laboratory

3.0 Responsibility/accountability:

Quality Control Chemist, Quality Assurance Officer, Quality Assurance Manager.

4.0 Abbreviations and Definitions

- cGMP – Current Good Manufacturing practice

- SOP – Standard operating Procedure

- PPE – Personal Protective Equipment

- Procedure

#### 5.1 Cleaning Procedure

- Turn off the flame photometer.

- Wipe the outer surface of the equipment using a clean, dry, lint-free cloth.

- Clean the nebulizer gently using cotton wool dampened with distilled water or ethanol.

- Ensure the work area around the instrument is free of spills and dust.

- Document your activity in the Flame Photometer Cleaning Logbook (BML-QDF-CL-012-F01)

#### 5.2 Operation Procedure

- Connect the equipment to the power source.

- Switch on the flame photometer using the ON/OFF button.

- Open the air compressor valve, then the fuel supply valve. Ensure all connections are secure.

- Press the ignition button to ignite the flame.

- Adjust the air pressure to 4.20 psi to ensure a stable blue flame.

- Aspirate Water for Injection (WFI) through the capillary tube for 15 minutes to clean and stabilize the system.

- Select the appropriate filter depending on the ion being measured (Na or K).

- Aspirate standard solutions (2 ppm, 4 ppm, 6 ppm, 8 ppm, 10 ppm) for calibration.

- Record readings and extrapolate the concentration of test samples using the calibration curve.

- Plot a graph of Concentration (ppm) vs Emission Intensity to quantify ion levels.

- After analysis, aspirate WFI again for 15 minutes to clean the system.

- Close the fuel valve and air compressor valve.

- Switch off the equipment from the ON/OFF button.

- Disconnect the equipment from the power supply.

- Document your activity in the Flame Photometer Usage Logbook (BML-QDF-CL-012-F02)

Materials, Equipment, Apparatus and Reagents

- PPE

6.0 REFERENCE/ENCLOSURE

In House

ENCLOSURE 1- FLAME PHOTOMETER CLEANING LOGBOOK- BML-QDF-CL-012-F01

ENCLOSURE 2- FLAME PHOTOMETER USAGE LOGBOOK- BML-QDF-CL-012-F02

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control head |
| --- | --- |
| Quality Control Analysts | Laboratory Assistants |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |