- OBJECTIVE:

To provide written guidelines for Cleaning and Operation of Stability Chamber

- SCOPE:

This SOP covers the process of Cleaning and Operation of Stability Chamber used in the chemical laboratory

- RESPONSIBILITY/ACCOUNTABILITY:

Quality Control Chemist, Quality Assurance Officer, Quality Assurance Manager.

4.0 ABBREVIATIONS AND DEFINITIONS:

-cGMP – Current Good Manufacturing practice

-SOP – Standard operating Procedure

-PPE – Personal Protective Equipment

- PROCEDURE:

Cleaning

| Activity | Frequency | Instructions |
| --- | --- | --- |
| Internal and external chamber surfaces | Monthly | Use clean, lint-free cloth and demineralized water or approved cleaning agent |
| Boiler tank, Heater and Reservoir | Weekly | Disconnect power; clean with approved cleaning agent, rinse thoroughly. |
| Clean the RH Sensor Area | Monthly | Gently wipe with demineralized water; avoid direct contact with sensor tip. |
| Logbook entry of cleaning activities | Each Occurrence | Document date, time, person responsible, and remarks in Equipment Cleaning Logbook (BML-QDF-CL-022-F01).. |

Operation

- Connect the stability chamber to the power.

- Turn on the main switch provided on the main control panel.

- Main screen displays “Fix Running”, the set print and process value of the controlling sensor.

  - H-display indicate that Humidity Sensor is running.

  - T-Display indicate that temperature sensor is running

  - Just below the sensor status, safety and alarm controller is in display

- Set the temperature (400C) on the temperature controller by pressing increment and decrement key.

- Set the humidity (75%RH) on the humidity controller by pressing increment and decrement key.

- Set the safety controller and the required alarm band and initial time. Safety controller will cut off the main contactor in case of overshoot or undershoot of the temperature is above the alarm band.

- The blower, air heater and humidity system will start after three (3) minutes fifty (50) seconds.

- The chamber light should be switched ON only for the observation purpose or to illuminate the internal chamber. Normally the chamber light should be kept OFF.

- Take printout of the stability temperature / humidity record data periodically (3 -4 times in a month) and critically review the data and make comment

- Take a back-up of the recorded electronic date every six (6) months and preserve/store properly in the document archive.

- Daily check the stability chamber. In case there is any anomaly found in the stability chamber. e.g. temperature and humidity settings, immediately call for maintenance. If the problem is rectified within short period, then document the anomaly situation. In the event that the rectification takes longer time, then transfer the product to another stability chamber unit with similar stability study condition until the stability chamber is rectified and raises a deviation note, Corrective and Preventive Action (CAPA) and Equipment Maintenance Request.

- Record the activity in the usage logbook (BML-QDF-CL-022-F02).

Materials, Equipment, Apparatus and Reagents

- PPE, demineralized water

6.0 REFERENCE/ENCLOSURE:

- Equipment manual

- ENCLOSURE 1- Equipment Cleaning Logbook- BML-QDF-CL-022-F01

- ENCLOSURE 1- Equipment Usage Logbook- BML-QDF-CL-022-F02

7.0 TRAINING INVOLVEMENT:

| Quality Assurance Manager | Quality Control head |
| --- | --- |
| Quality Control Analysts | Laboratory Assistants |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY:

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |