- OBJECTIVE:

To describe the procedure for offline analysis of Total Organic Carbon (TOC) in samples using a UV photo-oxidation TOC analyzer.

- SCOPE:

This SOP applies to QA/QC personnel analyzing water, pharmaceutical samples for TOC using the offline method.

- RESPONSIBILITY/ACCOUNTABILITY:

- Analyst: To carry out the TOC analysis and record results.

- QA/QC Supervisor: To ensure compliance with the SOP and review results.

4.0 ABBREVIATIONS AND DEFINITIONS:

- TOC: Total Organic Carbon.

- UV Photo-Oxidation: Oxidation of organic carbon using UV light.

- PROCEDURE

## Instrument Preparation

## 1. Power on the TOC analyzer and allow it to warm up (typically 60 minutes).

## 2. Ensure the water bottles are filled with the correct solutions.

## 3. Check waste bottle and empty if nearing capacity.

## 4. Verify that UV lamp status is OK.

## Sample Preparation

1. Rinse sample glassware 3× with high-purity water.

2. Collect sample into TOC glassware.

3. Cap tightly and label clearly with sample ID.

## Measurement (Offline Mode)

1. Enter your unique password

2. Login

3. Enter the username and company name

4. Click on OK

5. Select offline mode from the software, then click OK

6. Put the aspirating tube into the sample container.

7. Use the sample to flush the system through the rinse cycle.

8. Click on measure, then on the measure setting.

9. Give your sample an I.D name or number

10. Click on finish & save

11. Click on START

12. Record TOC value (in µg/L).

## Replicates and Controls

## Run each sample in triplicate.

## Include blank (high-purity water) and control sample (known TOC) with every batch.

# Post-Run Activities

1. Rinse lines and container using high-purity water.

2. Put off instrument when analysis is completed.

3. Dispose of waste.

4. Record all results in its sample record.

## Periodic Calibration

1. Prepare calibration standards (e.g., 0, 1, 2, 5 ppm TOC).

2. Run calibration sequence.

3. Validate calibration curve (R² ≥ 0.995).

# Troubleshooting

| Issue | Action |
| --- | --- |
| High blank reading | Check water purity, clean tube |
| Irregular TOC values | Check for bubbles, verify using high purity water. |
| Carryover contamination | Perform deep cleaning or rinse loop |

# Documentation

- TOC CLEANING LOGBOOK

- TOC USAGE LOGBOOK

- TOC CALIBRATION LOGBOOK

# Materials and Equipment

- UV Photo-Oxidation TOC Analyzer (BIOBASE)

- Sample tubes (pre-cleaned)

- High-purity water (TOC ≤ 100 ppb)

- Pre-treated glassware

- Waste container

- PPE (gloves, lab coat, goggles)

- REFERENCE/ENCLOSURE

INSTRUMENT USER MANUAL

USP <643> AND <645>

ICH Q2(R1) FOR ANALYTICAL METHOD VALIDATION

ENLCOSURE 1- TOC CLEANING LOGBOOK- BML-QDF-CL-028-F01

ENCLOSURE 2- TOC USAGE LOGBOOK- BML-QDF-CL-028-F02

ENCLOSURE 3- TOC CALIBRATION LOGBOOK- BML-QDF-CL-028-F03

7.0 TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control head |
| --- | --- |
| Quality Control Analysts | Laboratory Assistants |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE/ DOCUMENTATION UNIT |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |