- Objective:

To provide written guidelines for cleaning and Operation of Electronic weighing balance.

2.0   Scope:

This SOP covers the process of cleaning and Operation of Electronic weighing balance.

3.0 Responsibility/accountability:

Laboratory Attendant, Quality Control Microbiologist, Quality control Head, Quality Assurance Manager, Superintendent Pharmacist.

4.0 Abbreviations and Definitions

- IPA: Isopropyl Alcohol

- PPE: Personnel protective Equipment

- Procedure

Safety Precaution

- Ensure the Weighing balance is switched OFF before cleaning.

- Check the main power source and ensure it is OFF

- Do not MOVE or change the position of the weighing balance while cleaning

- Do not place wet or hot materials directly on the weighing balance

Cleaning Procedure

Cleaning should be performed before and after use

Before Use

- Ensure the balance is switched OFF before cleaning

- Open the side glass and clean the top load and inside cubicle with a dry lint-free cloth before use and at least twice a week when not used so that there will be no dust particles on the weighing balance

- Clean the top glass and side glass with dry cloth.

- Re-clean the inside cubicle and top load with 70%IPA moist cloth and wipe dry after cleaning

After Use

- Switch OFF the weighing balance

- Remove any material from the balance

- Clean the balance with a lint-free cloth and 70% IPA

- Record the activity in the equipment cleaning Logbook (BML.QDF.MC.005-F01).

Operation Procedure

- Ensure the cleanliness of the instrument

- Ensure the weighing balance is placed on a stable and horizontal surface

- Ensure the balance is connected to power

- Switch on the balance from the adaptor power unit and press the ON/OFF button on the balance panel to power it on

- Ensure the display reads zero

- Allow the weighing balance to warm up for 15 minutes

- Calibrate the weighing balance by following the procedure in the calibration SOP. BML.QDF.MC.045

- Place foil paper on the pan and tare to zero the balance

- Place materials to be weighed on the foil paper with a clean spatula up to the desired target weight

- After using the weighing balance ensure it’s clean and not soiled by any media or materials.

- All weighing shall be recorded in the usage log book (LGB-QDF-MC-005).

Materials, Equipment, Apparatus and Reagents

- PPE

- Foil paper

- Spatula

- Cotton wool

- Lint-free Cloth

- 70% IPA

- Hand Gloves

- REFERENCE/ENCLOSURE

REFERENCE: EQUIPMENT MANUAL

ENCLOSURE 1- ELECTRONIC WEIGHING BALANCE CLEANING LOGBOOK- (BML.QDF.MC.005-F01).

ENCLOSURE 2- ELECTRONIC WEIGHING BALANCE USAGE LOGBOOK- (LGB-QDF-MC-005).

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality Control Head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |