- Objective:

To provide written guidelines for cleaning and operation of Bucket Autoclave

- Scope:

This SOP covers the process of cleaning and Operation of Bucket Autoclave used in Microbiology Laboratory.

3.0 Responsibility/accountability:

Laboratory Assistant, Quality Control Microbiologist, Quality Assurance Manager, Supt. Pharmacist.

4.0 Abbreviations and Definitions

N/A

- Procedure

Cleaning Procedure:

- Switch off the autoclave by pressing the red button.

- Clean the bucket autoclave with soap solution

- Rinse thoroughly with distilled water until the soap solution is removed.

- Record the cleaning in ‘equipment cleaning record’ (BML-QDF-MC-007-F01).

Operation procedure:

For the sterilization of the following:-

- Laboratory equipment (Glass and metal)

- Culture Media

- Pharmaceutical operations in sealed containers

- Any other suitable materials.

# OPERATIONS:

- Ensure that there is sufficient water in the outer chamber of the sterilizer (About 2 - 3.5 liters of water is sufficient).

- Place articles to be sterilized in the inner container. The arrangement of materials to be sterilized should allow free movement of vapor.

- Place rightly the cover. Locate the six-drop bolt and tighten the six-drop bolts and wing nuts in opposite pairs until the lid is sealed on the pot. (Normal finger – tighten is quit adequate).

- Put on the plug and switch on the main electrical supply.

- Move both switches at the base of the sterilizer to high position (down wards).

- Allow the gauge to register 1040C as the water boils and then close the black knob (Air cock) on top of the lid.

- The gauge commences to register pressure and when the temperature reached 1210C, It Is only now that sterilization begins. The sterilization time is 20minutes for equipment and preparations and 15minutes for culture media.

- Turn off the switch marked off and allow the temperature to drop from 121 - 900c, Open the air cock to release pressure from the apparatus.

- Disconnect from the main electrical supply.

- Leave the autoclave to naturally cool down until the gauge reads zero.

- Record the usage in ‘equipment usage record’ (LGB-QDF-MC-007)

LID OPENING PROCEDURE

- Check that the gauge is at zero.

- Unscrew the six wing nuts and then two collared wing nuts to full turn. This allows the lid to rise to a safe position while it is still being held captive, releasing any residual pressure.

- Finally, fully unscrew the six nuts.

# CAUTION

- Before opening the lid, the vacuum must be broken by opening air cock (Black knob on the lid).

- Ensure that the sterilizer is fully vented

Materials, Equipment, Apparatus and Reagents

- PPE

- Foil paper

- Distilled water

Prepared media

- REFERENCE/ENCLOSURE

REFERENCE: EQUIPMENT MANUAL

ENCLOSURE 1- BUCKET AUTOCLAVE CLEANING LOGBOOK- (BML-QDF-MC-007-F01).

ENCLOSURE 2- BUCKET AUTOCLAVE USAGE LOGBOOK- (LGB-QDF-MC-007).

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |