- Objective:

To provide written instruction on sterility testing on finished product

- Scope:

To provide written guidelines for carrying out sterility tests on finished products.

3.0 Responsibility/accountability:

Quality control head, Quality assurance manager and superintended pharmacist

4.0 Abbreviations and Definitions

LAF – Laminar Airflow

TSB – Tryptone soy broth

FTM - Fluid Thioglycollate medium

DHS -   Dry heat sterilization

- Procedure

- Follow SOP, for entry and exit procedure to microbiology laboratory (BML.QDF.MC.001).

- Ensure the sterile room is cleaned, the workbench and the laminar air flow is disinfected with 70% IPA as per SOP for cleaning micro lab and SOP for cleaning work bench

- Sterilize all required accessories for sterility test in autoclave at 1210c for 20 minutes as per sop for sterilization by autoclave and glassware by DHS at 110oc to 125oc as per sop for cleaning and sterilization of glassware

- ON the UV light and LAF before 30 minute from start testing

- Maintain high level of precaution during test to avoid cross contamination

- Prepare TSB for fungal growth and FTM for bacterial growth based on the number of batches to be analyzed as per the SOP for preparation of culture media (BML.QDF.MC.011)

- Dispense 100ml of each broth into sterile bottles in triplicate for each batch and label appropriately with the product name, B.No. Date of testing, completion date and tested by. Sterilize for 20minutes at 121˚c and 15 psi pressure, allow cooling to about 45˚c

- For each batch produced, take twenty (20) samples (pouches) randomly, wipe individually with 70% IPA solution, and keep in a clean transparent bucket. Then transfer the sample into sterility room

- Arrange all the sterilized broth on the LAF before starting the test.

- Under the lamina flow, Cut the tip of the pouch with sterile SS scissor and aseptically pour the content in sterile conical flask. similarly, transfer the content of all 20 pouches to the sterile conical flask and mix

- Transfer 10ml sample from the conical flask into each of the triplicate 100ml sterile FTM bottles aseptically. Similarly, transfer 10ml sample into each of the triplicate 100ml sterile TSB bottles

- Simultaneously prepare a negative control by transferring 10ml sterile WFI to 100ml of sterile FTM and TSB bottles as a negative control

- Incubate the inoculated bottles i.e TSB at 20-25˚c and FTM at 30-35˚c for fourteen (14) days.

Observation and Interpretation of results

- Visually examine the media bottles daily to its conclusion for macroscopic evidence of microbial growth. If the sample being tested renders the medium turbid so that the presence or absence of microbial growth cannot be readily determined by visual examination before 14 days, transfer portions(each not less than 1ml) of the medium to fresh sample bottles of the same medium ,and then incubate the original and transferred sample bottles for not less than 4days.

- If no evidence of microbial growth is found, the product to be examined comply with the test for sterility

- If evidence of microbial growth is found, the product to be examined does not comply with the test for sterility, unless it can be clearly demonstrated that the test was invalid for causes unrelated to the product to be examined. The test may be considered invalid only if one or more of the following conditions are fulfilled:

- The data of the microbiological monitoring of the sterility testing facility show a fault

- A review of the testing procedure used during the test in question reveals a fault

- Microbial growth is found in the negative controls

- If the test is declared invalid, it is repeated with the same number of units as in the original test.

- If no evidence of microbial growth is found in the repeat test, the product examined complies with the test for sterility.

- If microbial growth is found in the repeat test, the product examined does not comply with the test for sterility.

- Record the activity/values of the finished products Sterility results in the various products forms.

| S/no | Product | Sterility Test Result Form No. |
| --- | --- | --- |
| 1 | Darrows full solution | BML-QA-STP-FPF-008-F05 |
| 2 | Isoplasma 4% solution | BML-QA-STP-FPF-014-F05 |
| 3 | Normal Saline (500ml) | BML-QA-STP-FPF-001-F05 |
| 4 | Dextrose 5% solution (500ml) | BML-QA-STP-FPF-003-F05 |
| 5 | Dextrose 5% Saline solution (500ml) | BML-QA-STP-FPF-002-F05 |
| 6 | Dextrose 4.3% saline | BML-QA-STP-FPF-012-F05 |
| 7 | Dextrose 10% solution | BML-QA-STP-FPF-013-F05 |
| 8 | Darrows half solution | BML-QA-STP-FPF-007-F05 |
| 9 | Mannitol 10% | BML-QA-STP-FPF-010-F05 |
| 10 | Mannitol 20% | BML-QA-STP-FPF-011-F05 |
| 11 | Ringers solution | BML-QA-STP-FPF-009-F05 |
| 12 | Biogyl solution | BML-QA-STP-FPF-015-F05 |
| 13 | Dextrose 50% solution | BML-QA-STP-FPF-018-F05 |
| 14 | Ciprofloxacin solution | BML-QA-STP-FPF-016-F05 |
| 15 | CAPD solution (1.5%) | BML-QA-STP-FPF-019-F05 |
| 16 | Fluconazole solution | BML-QA-STP-FPF-017-F05 |
| 17 | Normal Saline Solution (1000ml) | BML-QA-STP-FPF-004-F05 |
| 18 | Dextrose 5% solution (1000ml) | BML-QA-STP-FPF-006-F05 |
| 19 | Dextrose 5% Saline solution (1000ml) | BML-QA-STP-FPF-005-F05 |
| 20 | CAPD solution (1.5%) | BML-QA-STP-FPF-020-F05 |

Material and Equipment:

- Culture media (FTM and TSB)

- Portable Autoclave

- Incubator

- Hot air oven

- Sterile McCartney bottles

- LAF Cabinet

- Conical Flask and Beakers

- Measuring Cylinder

- Weighing Balance

- Distilled Water

- Foil Paper

- Paper Tape

- 70% IPA

- Surgical scissors

6.0 REFERENCE/ENCLOSURE

- STERILITY TEST RESULT FORMS- BML-QA-STP-FPF-(001…..020)-F05

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |