- Objective:

To provide written guidelines for cleaning, operation and calibration of heating block

- Scope:

This SOP covers the process of cleaning, Operation and calibration of heating block for          microbiology laboratory

- Responsibility/accountability:

Microbiologist.

4.0 Abbreviations and Definitions

N/A

- Procedure

- CLEANING:

- clean the external surface of the heating block with lint free wet cloth moist with 70%IPA

Frequency: As and when required

- Record the activity in the Heating Block Cleaning Record Book (BML-QDF-MC-035-F01).

- OPERATION:

- Connect the power cord of the heating block to the electrical main switch

- Switch ON mains and switch ON the power switch present on front panel of the heating block.

- Set the ‘heater capacity’ knob at 37°c (±1°c)

- Wait for the temperature to reach the set temperature and stabilize

- Place the assay tubes into the wells for incubation for the specified time (1 hr)

- After the incubation, remove the test tubes and switch OFF the instrument by pressing the ‘power’ switch

- Document the operation int the Heating Block Usage Logbook (LGB-QDF-MC-035)

- CALIBRATION:

  - Set the required temperature that is 37°c for calibration

  - Take 5 endotoxin test tubes (10 x75mm)

  - Add approximately 0.5ml of purified water in each tube for performing the calibration

  - Place the test tubes in the wells for few minutes to stabilize the water in the tubes

  - Place a calibrated thermometer in one of the tube for few seconds

  - Record the temperature and note down the display temperature of heating block

  - After note down the reading, take out and wipe the thermometer with the help of a cotton wool

  - Continue the process for all the remaining four tubes of LAL heating block

  - Record the activity in the Heating Block Calibration Form (BML-QDF-MC-035-F02).

INTERNAL: calibration shall be done quarterly

EXTERNAL: shall be done once in a year

Materials, Equipment, Apparatus and Reagents

- PPE

- Endotoxin Test tubes (10x75mm)

- Calibrated Thermometer

- 70% IPA

- Soft napkin

- Cotton wool

6.0 REFERENCE/ENCLOSURE

ENCLOSURE 1- HEATING BLOCK CLEANING FORM- (BML-QDF-MC-035-F01).

ENCLOSURE 2- HEATING BLOCK USAGE LOGBOOK- (LGB-QDF-MC-035).

ENCLOSURE 1- HEATING BLOCK CALIBRATION FORM- (BML-QDF-MC-035-F02).

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | MICROBIOLOGY LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |