- Objective:

To provide guidelines for Operation and calibration of Micropipettes

- Scope:

This SOP covers the process of Operation and calibration of micropipette used in Microbiology Laboratory.

- Responsibility/accountability:

Laboratory Attendant, Quality Control Microbiologist, Quality Assurance Manager.

4.0 Abbreviations and Definitions

µl   =    micro liter

mm  =  millimeter

- Procedure

Volume setting

The volume can be changed continuously by rotating the setting of the ring

Aspirating liquid

- The liquid, which is to be aspirated, is taken from a suitable vessel

- Attach the tips to the pipette firmly

- Immerse the pipette tips approximately 3mm to 5mm into the liquid or suitable depth into the liquid

- Allow the control button to slide back slowly

- Pull the micropipette tips out of the liquid slowly

Dispensing liquid

- Hold the tips at an angle against the inside wall of the suitable vessel

- Press down the control button slowly to the first stop and wait until the liquid stops flowing

- Press down the control button to the second stroke until the tip is completely empty

- Hold down the control button and pull the tips up the inner wall of the tube

- Allow the control button to slide back slowly

CALIBRATION

- Calibrate the micropipettes in a vibration free room at a constant temperature of water between 20 to 250c.

- Calibrate the micropipettes as per given in table -1.

- First, determine the density factor of water and then set desired volume using the volume-setting knob

- Carefully fit tip into tip cone and flush with distilled water by pipetting the adjusted volume at least 3-4 times

- Carefully aspirate the liquid and pipette the sucked volume of distilled water into a tarred volumetric flask

- Repeat at least five times and record the result

- Calculate the pipette volume as per following formula below:

Volume = weight of water sample/ specific gravity of water (at 250c, 0.9982)

- The percent variation can be calculated as formula given below:

% variation = 100 ₓ actual weight – observed weight

_________________________

Actual weight

- Precaution: use the analytical balance with a readability of 0.01g

Frequency: Quarterly

Table-1

| Micropipettes ID | Volume taken in µl | Tolerance Limit in µl |  |
| --- | --- | --- | --- |
|  |  | min | max |
| 1 | 100 | 99.9 | 100.1 |
|  | 500 | 499.9 | 500.5 |
|  | 1000 | 999.0 | 1001.0 |
| 2 | 10 | 09.99 | 10.01 |
|  | 50 | 49.95 | 50.05 |
|  | 100 | 99.9 | 100.1 |

Materials, Equipment, Apparatus and Reagents

- PPE

- Micro pipette

- Micro tips

- Conical flask

- Weighing balance

- Distilled water

- Calibrated thermometer

Precautions:

- Do not lay down the pipette when a filled pipette tip is attached

- Do not allow any liquid to enter into the pipette

- Do not clean the pipette with acetone or aggressive solutions

- REFERENCE/ENCLOSURE

INSTRUCTION MANUAL

ENCLOSURE 1- MICROPIPETTE CLEANING FORM- (BML-QDF-MC-039-F01).

ENCLOSURE 2- MICROPIPETTE USAGE LOGBOOK- (LGB-QDF-MC-039).

ENCLOSURE 1- MICROPIPETTE CALIBRATION FORM- (BML-QDF-MC-039-F02).

- TRAINING INVOLVEMENT

| Quality Assurance Manager | Quality control head |
| --- | --- |
| Quality Microbiology staff | Laboratory assistant |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | MICROBIOLOGY LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 05 | ### 04 | To Update |
| ### 06 | ### 05 | - Inclusion of enclosures format - Change in sop format generally |