- OBJECTIVE:

This SOP describes the process of Handling and operation of Hanna pH Meter

- SCOPE:

This SOP covers the usage and operation of the Hanna pH Meter

- RESPONSIBILITY/ACCOUNTABILITY:

- The Chemical analyst shall calibrate the pH meter using appropriate buffer solutions.

- The user must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS:

- PPE - Personnel Protective Equipment

- pH meter: A scientific instrument that measures the hydrogen ion activity in water-based solutions, indicating its acidity or alkalinity expressed as pH.

- PROCEDURE:

GENERAL:

The Hanna microprocessor-based pH meter is supplied complete with a glass body combination electrode, a temperature probe, an electrode holder and 20ml each of pH 4.01 and 7.01 buffer solutions, an electrolyte solution, a 12DC power adapter and an instructional manual.

The equipment is assembled according to the instructional manual and connects the system to a power source.

Calibration

- The meter comes factory calibrated for temperature so calibration is only required for pH calibration should be carried out often, at least once a week. Re calibration is also recommended for higher accuracy, when the electrode is replaced, when testing aggressive chemicals and after the Reset button is pressed.

- To calibrate the meter pour small quantities of both buffers into separate clean beakers leaving aside an amount of either buffer for cleaning the probe after each calibration.

- Immerse the pH electrode and temperature probe approximately 4cm into the buffer solution chosen. Keep both electrode and probe close within the solution

- Press CAL. The CAL and BUF indicators and the most common 7.01 buffer will be displayed on the secondary LCD.

- Using the UP or Down keys select different buffer values

- The NOT READY indicators will blink on the LCD until the reading has stabilized.

- When the reading is stable and close to the selected buffer the READY indication will appear on the LCD screen and CFM will blink. Press CFM to confirm the calibration

- The calibrated value is then displayed on the primary LCD and the secondary LCD will display the second expected buffer value

- For a one-point calibration, using only one buffer press CAL to quit calibration. For a two-point calibration using two buffers do not press the CAL button

- Rinse the probe and electrode in one of the beakers containing the second buffer to be calibrated

- Shake off the excess buffer and repeat steps, the machine is now calibrated and ready for use

Sample Measurement

  - Ensure that the machine has been calibrated and the electrode is in working order.

  - The solution to be tested should be placed in two clean beakers, one for measurement and the other for rinsing the electrode prior to testing

  - Do not switch on the machine until both the electrode and the temperature probes are immersed 4cm into the test solution

  - Once the machine is switched on , it will adjust for  temperature and the microprocessor will attempt to correctly identify the PH, therefore the reading will take few minutes to be ascertain

  - Stir the solution around the probe and the electrode. The PH value should be recorded once the reading on the LCD screen has stabilized

  - Remove the probes from the solution and rinse with cleaning solution or distilled water

  - Insert the electrode in to the protective plastic cap containing the storage solution (HI70300). And return the electrode.

  - Record the usage activities in the user logbook LGB/QCS/CL/008/01 for the pH meter appropriately.

pH meter User Logbook LGB/QCS/CL/008/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

  - NB: The pH meter electrode should under any circumstance never be tapped against any surface to avoid cracks or breakage.

Trouble Shooting

  - Noise: This causes fluctuating readings. Refer to manual for cleaning procedure or replace electrolyte within glass chamber of electrode

  - Dry membrane/junction: soak in storage solution for 1 hour.

  - Drifting:  Soak in warm filling solution (50-60OC) for 1 hour and then rinse with distilled water. Refill electrode chamber with filling solution

  - Error messages: Errors messages for wrong buffer might require a buffer change, persistent error messages even after this and cleaning procedures mean the electrode is faulty and needs replacement.

Materials, Reagents, Equipment, Apparatus

- PH 4 Buffer

- PH 7 Buffer

- Glass cylinders

- De-ionsed water

- REFERENCE/ENCLOSURE

- Hanna pH instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Chemical Analyst |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |