- OBJECTIVE:

To provide written guidelines for cleaning and operation of the Moisture Content Analyzer

- SCOPE:

This SOP covers the procedure for cleaning and operating the Moisture Content Analyzer

- RESPONSIBILITY/ACCOUNTABILITY:

- The user must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

Moisture content is the relative water content in granules, measured in percentage.

- PROCEDURE

Cleaning

- Disconnect from the power source.

- Open the cover and remove the pan handler, sample pan, draft shield, and heat shield from the instrument before cleaning.

- Use a lint-free cloth for cleaning and clean the exterior of the instrument and the drying compartment components with a mild cleaning agent. If the inside of the glass is dirty, lift up the heating chamber, release the glass for cleaning by pressing both sides of the glass holder.

- Never use abrasive cleaning agents or solvents!

- Ensure that no liquid enters the interior of the instrument.

- Replace and reassemble the components after cleaning.

Operation

- Connect the power cord to the power supply socket

- Allow the moisture analyzer to warm up for at least 30minutes Press the                  button to switch the Moisture Analyzer ON.

- Open the cover, and clean the pan.

- Place the pan handler in the sample chamber, ensure that the tongue of the pan handler fits exactly in the slot of the draft shield element.

- Place the empty sample pan and glass fiber pan in the pan handler. The sample pan must lie flat in the pan handler

- Close the cover, and this sets the Moisture Analyzer to zero automatically.

- Open the cover on the Moisture Analyzer.

- Put sample into the glass fiber pan NMT  0.5 to 1 gram of sample

- Close the cover, the Moisture Analyzer starts the drying and measurement process automatically

- Follow the progress of the drying cycle on the display.

- There are two displays available when the Moisture Analyzer is running a test. One showing the %MC value and another showing a curve.

- Press the %MC the display will switch to show curve

Attention: Minimum sample size must be greater than 0.5 gram.

- Switch off after use

CAUTION

- To avoid electric Shock Hazard, Disconnect the equipment from the power supply before cleaning.

- During operation, never open the heating module itself as the ring-shaped heating reflector or its protective glass can reach 400 ° C!

- Do not touch or clean the halogen heater.

- Do not place any combustible materials on, under or next to the instrument n ,>

- The vent over the sample must never be covered, plugged, taped over or tampered with in any other way

- Never leave the instrument unattended!

Materials, Equipment, Apparatus and Reagents

- PPE’s

- Soap

- Demineralized water

- REFERENCE/ENCLOSURE

- Ohaus moisture analyzer instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Chemical Analyst |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |