- OBJECTIVE:

To provide written guidelines for cleaning and Operation of Microbiology Laboratory Oven

- SCOPE:

This SOP covers the process of cleaning and Operation of Microbiology Laboratory Oven

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

PPE -   Personal protective equipment

- PROCEDURE

Cleaning Procedure:

- Switch OFF the electric oven

- Disconnect the oven from the power source.

- Allow the oven to cool to room temperature.

- Wipe all the internal surface, walls, top, bottom and trays of the oven using a lint-free cloth dampened with mild detergent solution.

- Wipe again using a cloth

- Wipe the surface, walls, top, bottom and trays of the oven with dry lint free cloth soaked with 70%IPA on daily basis so that there will be no dust particles in the oven.

- Wrapped all the glassware to be used for microbiological analysis with foil paper.

- Lift the apparatus from its original position and clean the base.

- Return the apparatus /glassware to its original position.

Operation Procedure:

- Ensure the cleanliness of the instrument

- Open the ventilation knob provided on top of the oven

- Connect the switch to main electrical supply.

- Set the thermostat to the required temperature.

- Select a medium of drying/sterilization.

- Ensure that glass wares are properly arranged in the oven, when drying/sterilizing.

- Ensure that the door is well closed when in use.

- Upon completion of the sterilization time, fill in the oven user logbook provided.

Oven User Logbook LGB/QDS/MC/006/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- Switch off and allow cooling down naturally before the door is opened.

- Disconnect the cable from the main electricity supply.

Materials, Equipment, Apparatus and Reagents

- PPE

- Foil paper

- Cotton wool

- 70% IPA

- Glass wares

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

    - Medfield oven instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |