- OBJECTIVE:

To provide written guidelines for cleaning and operation of Bucket Autoclave in microbiology laboratory of Biomedical Limited.

- SCOPE:

This SOP covers the process of cleaning and operation of Bucket Autoclave in microbiology laboratory

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist shall strictly follow this procedure to achieve safe and efficient use of equipment

- The user shall ensure the equipment is kept clean and reports any malfunctioning.

- The Quality Control Head ensures scheduled maintenance is carried out.

4.0 ABBREVIATIONS AND DEFINITIONS

- Nil

- PROCEDURE

Cleaning Procedure:

- Switch off the autoclave by pressing the switch button.

- Clean the bucket autoclave with soap solution

- Rinse thoroughly with distilled water until the soap solution is removed.

- Record the cleaning of the equipment cleaning record.

Operating procedure for sterilization:

- Laboratory equipment (Glass wares and metals)

- Culture Media

- Pharmaceutical operations in sealed containers

- Any other suitable materials.

# Operations:

- Ensure that there is sufficient water in the outer chamber of the sterilizer (About 2 - 2.5 liters of water is sufficient).

- Place articles to be sterilized in the inner container. The arrangement of materials to be sterilized should allow free movement of vapor.

- Place rightly the cover. Locate the six-drop bolt and tighten the six-drop bolts and wing nuts in opposite pairs until the lid is sealed on the pot. (Normal finger – tighten is quit adequate).

- Put on the plug and switch on the main electrical supply.

- Allow the gauge to register 1040C as the water boils and then close both safety valve and release valve (Air lock) on top of the lid.

- The gauge commences to register pressure and when the temperature reached 1210C, It Is only now that sterilization begins (release of pressure).

- The sterilization time for culture media and glass wares is 15 minutes and decontamination time is 15-30 minutes.

- Turn off the switch and allow the gauge indicator to drop down to zero.

- Open the valves to release pressure from the apparatus.

- Disconnect from the main electrical supply.

- Leave the autoclave to naturally cool down.

- Log in all activities in the Autoclave user’s logbook (LGB/QDS/MC/007/01)

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

Lid Opening Procedure

- Check that the autoclave is cooled down to a desired temperature that bare hands could withstand.

- Unscrew the six wing nuts, this allows the lid to rise to a safe position while it is still being held captive, releasing any residual pressure.

- Finally, fully unscrew the six nuts.

# Note:

Before opening the lid, the vacuum must be broken by opening air cock (pressure gauge on the lid). Ensure that the sterilizer is fully vented.

Materials, Equipment, Apparatus and Reagents

- PPE

- Foil paper

- Distilled water

- Prepared media

- Glass wares

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

- Bucket autoclave instruction manual.

- TRAINING INVOLVEMENT

| Quality Control Head | Microbiologist, Laboratory assistance |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |