- OBJECTIVE:

To describe the process for cleaning, operating and calibrating the incubator to ensure the incubator provides the heat required (Temperature) for optimum growth of microorganisms

- SCOPE:

This SOP covers the process of cleaning and Operation of incubators.

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

- Nil

- PROCEDURE

Cleaning Procedure:

    - Switch off the incubator.

      - Remove materials from the incubator to be cleaned to another incubator at the same temperature.

    - Clean the inner area with a disinfectant 70 % ethanol solution or 1% Oxonian solution 	using a lint-free cloth.

    - Make the area completely dry with lint-free cloth.

    - After complete drying, switch on the incubators allow temperature to attain the required 	temperature.

    - With the temperature confirmed, the incubator is ready for use.

Operation Procedure

- Connect the switch to main electrical supply.

- Connect the equipment to the main electrical supply

- Set the thermostat to the required temperatures (i.e 20-250C ± 2.5oC for fungi growth and 30-350C ± 2.5oC for Bacteria).

- Incubate the plates turn upside down except in special cases.

- Ensure the doors are well closed.

- After successful operation, switch off the appliance and disconnect the cable from the main electrical supply.

- Record the usage activities in the user logbook LGB/QDS/MC/009/01 for the incubator appropriately.

Incubator User Logbook LGB/QDS/MC/009/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- Clean the incubator with 70% IPA when the Incubator assumes room temperature.

Materials, Equipment, Apparatus and Reagents

  - PPE

  - Foil paper

  - Cotton wool

  - 70% IPA

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

- Equipment user manual [medfield]

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |