- Objective:

To provide written guidelines for calibration of Glassware.

- Scope:

This SOP covers the procedure for Calibration of Glassware.

- Responsibility/accountability:

Laboratory Assistant, Quality Control Chemist, Quality Control Microbiologist, Quality Assurance Manager, Supt. Pharmacist.

4.0 Abbreviations and Definitions

N/A

- Procedure

Calibration:

- Volumetric glassware’s are calibrated at 23° to 27°C.

- Clean the glassware for calibration to remove any grease, dirt etc. by keeping it immersed in acid/NaOH for 30 minutes.

- Rinse the glassware under running tap water then rinse it twice with demineralised water and dry it at 60-65°C for 60 minutes. Ensure that glassware is completely dry.

- Remove the glassware from the oven and keep it in desiccators, ensure that the ambient temperature is 25°C ± 2°C.

- Arrange a water bath consisting of plastic beaker contain water at 25°C ± 2°C.

- Maintain the temperature of water bath at 25°C ± 0.5°C throughout the calibration.

- Place the apparatus to be calibrated on the balance and note the weight (A), remove the apparatus and transfer the distilled water at 25°C so that the meniscus of liquid in the apparatus on the mark level. Weight the apparatus and note down the weight (B).

- Calculate the volume of glass ware using following formula i.e. (B – A)/ 0.99602

- Take the volumetric flask for calibration (10 to 1000 ml)

- Take graduated pipette for calibration (1 to 50 ml)

- Take Transfer pipette for calibration (25 ml)

- Take burette for calibration (50ml)

Materials, Equipment, Apparatus and Reagents

PPE

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

IN-HOUSE

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |