- Objective:

To provide written guidelines for Calibration of Stability chamber.

- Scope:

This SOP covers the procedure for Calibration of Stability chamber.

- Responsibility/accountability:

Laboratory Assistance, Quality Control Chemist, Quality Control Microbiologist, Quality Assurance Manager, Supt. Pharmacist.

4.0 Abbreviations and Definitions

N/A

- Procedure

Calibration

Temperature and Humidity performance

- Use a temperature and humidity sensor logger with a resolution not greater than 0.10C and a temperature range of 16-600C. The temperature sensor should be placed in the geometry of the effective space studio.

- Choose one point within the product temperature control range, set the SV temperature and humidity control value, when the PV measurement value is equal to the set value.

- Keep the temperature and humidity constant for about 1- 2hrs

- Observe the actual measured temperature and humidity value

- The difference between PV and the measured value displayed by the temperature controller should be ±1.50C and ±3.0%RH for humidity.

- Record your readings and observation

ALARM CHALLENGE TEST:

When a fault occurs and the temperature and/or humidity is out of control

- Observe the stability chamber over a period, through inconsistent water supply, to ascertain its alarm sensitivity.

PRECAUTION

Clean the humidifying tube to clean scale and other dirt regularly

Materials, Equipment, Apparatus and Reagents

Temperature and humidity sensor logger

Safety, Health and Environment Precautions

- Wear appropriate PPE

- REFERENCE/ENCLOSURE

Instrument manual

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |