- Objective:

To lay the procedure for the calibration of the pH meter.

- Scope:

This SOP covers the procedure for the calibration of the pH meter

- Responsibility/accountability:

Laboratory Assistants, QA/QC analyst, the Quality Assurance Manager and the Superintendent Pharmacist.

4.0 Abbreviations and Definitions

- GLP: Good Laboratory Practice

- GDP: Good Documentation Practice

- Procedure

- Plug the power line to a supply outlet and turn on the power switch

- Press “pH/mV” button make it into the pH correct measurement mode

- Press “TEMP” set the solution temperature (Temp indicator light flashes), then press “YES” button back to pH test status.

- Dip the electrode washed with demineralized water into standard buffer 7.00, the end of the electrode must be completely immersed into the buffer. Stir the electrode gently to create a homogeneous mixture.

- Wait for the measured pH value to stabilize press “STAND” (pH indicator light flashes), then press “YES” key to confirm calibration (pH indicator light stop flashing)

- The meter is now calibrated to the current buffer.

- Repeat the steps for additional calibrated point 4.00, 9.18

- Record the reading for the calibration.

- When all the calibration points set in the unit configuration set up are completed the meter returns to the pH measuring state

- Frequency for accuracy: calibrate every day, and as when required when drift is noticed.

If the measuring process is not normal, turn off the power and press “YES” key, then turn on the power to restore the initial state of the instrument and re-calibrate.

Storage

- Always keep the electrode bulb wet in pH 5.00 buffer or raw water.

- Do not store in deionised water

MATERIALS, EQUIPMENT, APPARATUS AND REAGENTS:

- Buffer

- Beaker

- Demineralized water

SAFETY HEALTH AND ENVIRONMENTAL PRECAUTIONS: PPE

- REFERENCE/ENCLOSURE

- Equipment Manual

- TRAINING INVOLVEMENT

| All departmental managers | All sectional heads |
| --- | --- |
| QA Doc. Unit staff |  |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| ### 00 | ### Nil | ### Nil | First |
| ### 01 | ### 00 | ### Jan 2024 | - Inclusion of enclosures format - Change in sop format generally |