- PURPOSE

  - To outline procedure for operation, calibration, verification and cleaning of Adam CCT4 4kg maximum capacity weighing scale at Emzor Pharmaceutical Industries Limited.

- SCOPE:

  - This procedure applies to operation, calibration, verification and cleaning of Adam weighing scale CCT 4 in Emzor Pharmaceutical Industries Limited, for its processes, and other necessary activities that have the potential to affect cGMP.

- RESPONSIBILITIES:

  - IPQA personnel are responsible for ensuring that this SOP is followed for the operation, calibration, verification and cleaning of the weighing scale.

- DEFINITIONS:

  - cGMP: Current Good Manufacturing Practices

  - QA: Quality Assurance

  - IPQA: In-Process Quality Assurance

- PROCEDURE:

  - OPERATION

    - Ensure that the pan is clean & free from dust.

    - Check and make sure that the scale is level indicator is at the centre of the circular marking and if necessary correct this by turning the levelling feet

5.1.3	  Connect plug to mains and switch ON the scale and printer.

5.1.4		The scale will turn on displaying the last weighing unit selected, either kilograms or pounds. To change the weighing unit press the [Units] button.

5.1.5		To change the weighing unit press the [SETUP] key and use the [1] or [6] buttons to scroll through the menu until ‘units’ appears on the display. Press [Tare] to select.

5.1.6		Press the [→0←] button to set the display to 0.0000.

5.1.7		To tare the scale, place a container on the scale and press [Tare].

5.1.8		As a product is added only the weight of the product will be displayed.

5.1.9		Press the [PRINT] button to print out displayed weight.

5.1.10	After weighing, remove all materials from atop the scale and press the [→0←] button to zero the scale.

5.1.11	Clean the scale immediately after use.

5.1.12	Sign the weighing scale usage logbook LGBK/QA/053.

| DATE | PRODUCT WEIGHED | BATCH NUMBER | TIME STARTED | TIME ENDED | DONE BY | CHECKED BY |
| --- | --- | --- | --- | --- | --- | --- |

5.2	CALIBRATION

5.2.1	Press [Tare] once, during the initial counting of the display after the power is turned on.

5.2.2	The “Count” display will show “P” requesting for the passcode number.

5.2.3	Enter the fixed passcode “1000”.

5.2.4	Press the [Tare] button.

5.2.5	The “Weight” display will show “u-CAL”

5.2.6	Press the [Tare] key and the “weight” display will show “noload” to request all weight be removed from the platform.

5.2.7	Press the [Tare] key to set the zero point.

5.2.8	The display will then show the calibration weight suggested in the “Count” display. If the calibration weight is different from the value shown, Press [CE] to clear the current value then enter the correct value as an integer value, it is not possible to have fractions of a kilogram or pound.

5.2.9	Press [Tare] to accept the calibration value and the “Weight” display will now show “Load”.

5.2.10	 Place the calibration weight on the platform and allow the scale to stabilize as indicated by the stable indicator.

5.2.11	Press [Tare] to Calibrate.

5.2.12	When calibration is done the scale will restart and return to normal weighing.

5.2.13	After calibration, the calibration should be verified using standard weights of 0.2kg, 1kg and 3kg.

5.2.14	Repeat calibration if deviation is outside ±0.0003kg.

5.2.15	Record displayed values of the weights in the verification and calibration logbook - LGBK/QA/054.

5.2.16	Affix Calibration label on the scale.

5.3	VERIFICATION

5.3.1	Verification shall be done daily.

5.3.2	Before verifying the scale, clean the weighing pan.

5.3.3	Ensure the level indicator is at the centre of the circular marking.

5.3.4	Switch ‘ON’ the main power supply.

5.3.5	Wait till the zero reading is displayed on the screen.

5.3.6	Allow to stabilize for 15 -20 seconds before placing standard weights on pan.

5.3.7	Place 0.2kg, 1kg and 3kg standard weights consecutively on the platform.

5.3.8	Allow to stabilize before taking reading.

5.3.9	Record displayed reading in the Verification and calibration of Adam weighing scale CCT4 logbook LGBK/QA/054.

5.3.10	Affix Verification label on the scale.

Daily verification and calibration logbook LGBK/QA/054

| Date | Standard weight (Kg) | Reading obtained (Kg) | Deviation (Kg) | Deviation (%) | Cal. Carried out | Cal. Weight (Kg) | Weight after Cal. (Kg) | Done by | Checked by |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

5.4	CLEANING

5.4.1	Cleaning should be carried out after every use or change of product.

5.4.2	Turn off scale and disconnect from mains.

5.4.3	Remove dust and powder using a dry napkin. Never blow; this could transport dirt or spilled product materials inside the scale.

5.4.4	For removal of sticky materials, use a damp lint-free cloth and mild solvent such as 70% ethanol or mild detergent.

5.4.5	Do not spray or pour liquids directly on the balance.

5.4.6	Remove parts for cleaning where possible e.g weighing platform.	 Only 	remove parts that can be removed without tools.

5.4.7	Clean the scale in its working location; do not tilt, move or carry it around if you are not instructed on how to carry a scale.

- REFERENCES

6.1 	Adam CCT scale operation manual.

  - Mettler Toledo: SOP for cleaning a laboratory balance https://www.mt.com/us/en/home/library/standard-operations-procedures/laboratory-weighing/Laboratory_Balance_Cleaning.html

7.0	CHANGE HISTORY:

| CURRENT REVISION | SUPERSEDE | REVIEW DATE | REASON FOR REVIEW |
| --- | --- | --- | --- |
| 00 | Nil | Nil | First issue |

- ENCLOSURE:

8.1      Not applicable

- DISTRIBUTION:

This is the various department/section where the SOP copies will be utilized.

| MASTER COPY | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPY | QUALITY ASSURANCE, PRODUCTION |