- OBJECTIVE:

To describe the procedure to ensures that the internal temperature and pressure readings are accurate and that the unit consistently achieves the necessary conditions for effective sterilization using biological and chemical indicator

- SCOPE:

To provide written procedure for calibration of Bucket Autoclave in the microbiological laboratory at both infusion and syrup plant in Biomedical Ltd.

- RESPONSIBILITIES/ACCOUNTABILITIES:

- Microbiologist: responsible for the operation of the instrument.

- Section in-charge: responsible for reviewing the procedure

- QC Head: responsible for implementing the procedure.

- QAM: overall compliance

- ABBREVIATION AND DEFINITIONS

QC- Quality Control

QAM- Quality Assurance Manager

- PROCEDURES

CALIBRATION:

- Ensure that the equipment is placed in appropriate place

- Before starting the cleaning procedure, disconnect the power supply to the Autoclave.

- Check the bolt of the Autoclave whether is intact

- Check the Air cock of the Autoclave

- Check the pressure gauge.

- Fill the bucket autoclave with water, make sure it cover the element

- Place chemical indicators on all items to be sterilized. Place Biological Indicator in "worst-case" or cold spots within a standard, properly loaded basket (e.g., top, middle, and bottom).

- Operate the autoclave through its normal sterilization cycle (e.g., 121°C at 15 PSI for 15-20 minutes).

- After the cycle and cooling, check the chemical indicators. The strip should have visible dark lines, indicating it reached the required temperature.

- Aseptically remove the BI vial and incubate the indicator at 55-60°C for 24hours

- Successful sterilization shows no color change, remains purple while a failed sterilization shows color change from purple to yellow

- Frequency: This process should be performed every 3 month or after major maintenance to ensure consistent results.

- Record all readings in the calibration logbook in the format below (BML.QA.MCB.AV.001)

Materials, Equipment, Apparatus and Reagents

- Chemical Strip

- Biological Indicator

- Bucket Autoclave

- REFERENCE AND ENCLOSURE

REFERENCE: EQUIPMENT MANUAL

ENCLOSURE 1- INCUBATOR CALIBRATION LOGBOOK - (BML.QA.MCB.AV.001)

7.0 TRAINING INVOLVEMENT

| Quality Control Manager | QC Microbiologist |
| --- | --- |
| QA Doc. Unit staff | Lab. Assistant |

8.0 DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | ALL DEPARTMENT |

9.0 CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in SOP format. - Inclusion of traceability of enclosures to parent SOP |