| S/N | TEST | DESCRIPTIONS OF PROCEDURE |
| --- | --- | --- |
|  | DESCRIPTION | Determine the colour and appearance of material |
|  | SOLUBILITY | Determine its solubility in water and in ethanol (96 per cent). |
|  | IDENTIFICATION | A.0.25 g gives the reaction of silicates. Mix the prescribed quantity of the substance to be examined in a lead or platinum crucible by means of a copper wire with about 10 mg of sodium fluoride R and a few drops of sulfuric acid R to give a thin slurry. Cover the crucible with a thin, transparent plate of plastic under which a drop of water R is suspended and warm gently. Within a short time a white ring is rapidly formed around the drop of water. B.1 mL of solution S neutralised with dilute sodium hydroxide solution R gives the reaction of magnesium. ### Solution S To 2.0 g add a mixture of 4 mL of nitric acid R and 4 mL of distilled water R. Heat to boiling with frequent shaking. Add 12 mL of distilled water R and allow to cool. Filter or centrifuge to obtain a clear solution and dilute to 20 mL with distilled water R. |
|  | ALKALINITY | To 10.0 g in a 200 mL conical flask, add 100.0 g of water R and heat on a water-bath for 30 min. Allow to cool and make up to the initial mass with water R. Allow to stand and filter or centrifuge until a clear liquid is obtained. To 10 mL of this liquid add 0.1 mL of phenolphthalein solution R. |
|  | WATER-SOLUBLE SALTS | In a platinum dish, evaporate to dryness on a water-bath 20.0 mL of the liquid obtained in the test for alkalinity. The residue, ignited to constant mass at 900 ± 50 °C, weighs a maximum of 30 mg. |
|  | CHLORIDE | Dilute 0.5 mL of solution S to 15 mL with water R. Prepare the standard using a mixture of 5 mL of chloride standard solution (5 ppm Cl) R and 10 mL of water R. |
|  | SULPHATES | Dilute 0.3 mL of solution S to 15 mL with distilled water R. |
|  | ARSENIC | Determined on 2.5 mL of solution S. |
|  | LOSS ON IGNITION | Determined on 0.5 g by ignition to constant mass at 900 ± 50 °C in a platinum crucible. |
|  | ACID ABSORBING CAPACITY | Suspend 0.25 g in 0.1 M hydrochloric acid, dilute to 100.0 mL with the same acid and allow to stand for 2 h in a water-bath at 37 ± 0.5 °C, with frequent shaking. Allow to cool. To 20.0 mL of the supernatant add 0.1 mL of bromophenol blue solution R and titrate with 0.1 M sodium hydroxide until a blue colour is obtained. |
|  | ASSAY | ### Magnesium oxide To 1.000 g in a 200 mL conical flask, add 35 mL of hydrochloric acid R and 60 mL of water R and heat in a water-bath for 15 min. Allow to cool, filter, wash the conical flask and the residue with water R and dilute the combined filtrate and washings to 250.0 mL with water R. Neutralise 50.0 mL of the solution with strong sodium hydroxide solution R (about 8 mL). Carry out the complexometric titration of magnesium (2.5.11). 1 mL of 0.1 M sodium edetate is equivalent to 4.030 mg of MgO. ### Silicon dioxide To 0.700 g add 10 mL of dilute sulfuric acid R and 10 mL of water R. Heat for 90 min on a water-bath with frequent shaking, replacing the evaporated water. Allow to cool and decant onto an ashless filter paper (diameter 7 cm). Wash the precipitate by decantation with 3 quantities, each of 5 mL, of hot water R, transfer it to the filter and wash it with hot water R until 1 mL of the filtrate remains clear after the addition of 0.05 mL of dilute hydrochloric acid R and 2 mL of barium chloride solution R1. Incinerate the filter and its contents in a platinum crucible, then ignite the residue (SiO2) at 900 ± 50 °C to constant mass. |
|  | MICROBIAL ENUMERATION TEST | Determination of total aerobic microbial count (TAMC) and total combined yeasts and molds count (TYMC) by the pour plate method. - Slightly open the lid of the petri plate and dispense 1 ml of diluted sample in the center of the Petri plate using a sterile calibrated pipette. - Open the lid of the prepared media. (soya bean digest agar for TAMC and Sabouraud Dextrose Agar for TYMC) - Pour about 15-20ml of the sterilized molten media at the appropriate temperature above the sample. - Close the lid of the plate then mix the sample in the media properly by gently swirling the plate. The plate is generally swirled in an “S” or “8” shape. - Allow the media to completely solidify. - Simultaneously perform the negative control by plating 1ml of diluent used instead of the sample for each medium. - Incubate the plate in an inverted position under suitable incubation conditions (For TAMC Incubate the plates of Soybean Casein Digest Agar at 30°C-35°C for 48hours and For TYMC incubate the plates of Sabouraud Dextrose Agar at 20°C - 25°C for 72hours Test for Absence of Pathogen Test for E. coli: Using same dilution - Transfer 1ml of the sample into a sterile McCartney bottle containing 9ml of sterile MacConkey broth. - Incubate the MacConkey broth bottle at 300C-350C for 24 hours. - After incubation shake the tube and transfer loop full portion of MacConkey broth on the surface of the pre incubated Eosin Methylene Blue Agar petri plate. - Incubate plates at 30°C to 35°C for 48 hours. - For negative control, proceeding same as above except without sample, streak a portion of incubated sterile MacConkey broth medium on the surface of pre incubated Eosin Methylene Blue Agar. - Observe and record the presence or absence of specified microorganisms. - Include positive and negative controls in each batch of tests. - Verify the performance of media and reagents used in the analysis. Interpretation of Results - The bacterial count should not exceed 102cfu/ml while that of the fungal count should not be more than 102cfu/ml. There should be absence of fecal coliform (E coli). - If growth occurs in the negative control plate, the test is invalid, failed negative control must be investigated - The total aerobic microbial count (TAMC) is considered to be equal to the number of CFU found using casein soya bean digest agar; if colonies of fungi are detected on this medium, they are counted as part of TAMC. - The total combined yeasts/mould count (TYMC) is considered to be equal to the number of CFU found using Sabouraud dextrose agar; if colonies of bacteria are detected on this medium, they are counted as part of TYMC - When an acceptance criterion for microbiological quality is prescribed it is interpreted as follows: 101 CFU : maximum acceptable count＝20 102 CFU: maximum acceptable count＝200, 103 CFU: maximum acceptable count＝2000. |

| - CURRENT REVISION | - SUPERSEDE | - REASON FOR REVIEW |
| --- | --- | --- |
| - 00 | - NIL | - FIRST ISSUE |
| - 01 | - 00 | - Review of the procedures - Change in header and footer format |