BIOMEDICAL LIMITED

CALIBRATION PROTOCOL/REPORT

BML/QA /CBP/001

| PROTOCOL NUMBER | BML/QA/CBP/001 |
| --- | --- |
| REPORT NUMBER | BML/QA/CBR/20/001 |
| VERSION NUMBER | 001 |
| EFFECTIVE DATE | APRIL 2020 |
| CALIBRATION DUE DATE | MARCH 2023 |

ADDRESS:                                               CONTACT DETAILS:

No. 1 Ohimege Road, Industrial Estate,        E-mail:

Ilorin, Kwara State,                                          Website:

Nigeria. 				            Tel.:

PROTOCOL APPROVAL PAGE

|  | WRITTEN: | REVIEWED: |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

|  | APPROVED | AUTHORIZED |
| --- | --- | --- |
| Signature: |  |  |
| Position: |  |  |
| Date: |  |  |

TABLE OF CONTENT

| S. No | Content | Page No. |
| --- | --- | --- |
| 1 | OBJECTIVE |  |
| 2 | SCOPE |  |
| 3 | CALIBRATION TEAM |  |
| 4 | PERIODIC RECALIBRATION |  |
| 5 | CALIBRATION |  |
| 6 | EQUIPMENT CALIBRATION |  |
| 7.0 | DEVIATION |  |
| 8.0 | SUMMARY/CONCLUSION |  |
| 9.0 | REPORT ENDORSEMENT |  |

- OBJECTIVE

The goal of this protocol is to show periodic calibration of the equipment QAQC controlled, which aims to minimise any measurement uncertainty by ensuring the accuracy of the test equipment. The calibration shall quantify and control errors or uncertainties within measurement processes to an acceptable level, thereby establishing confidence/reliability and traceability.

- SCOPE

This protocol shall be applicable to all QAQC equipment within.

- CALIBRATION TEAM

|  | Name | Signature/name |
| --- | --- | --- |
| Calibrated by |  |  |
| Verified by |  |  |
| Approved by |  |  |

- PERIODIC RECALIBRATION

In case there is a change or if the system is violated or as stipulated by the quality assurance system.

REASON FOR CALIBRATION:

The goal of calibration is to minimise any measurement uncertainty by ensuring the accuracy of test equipment. I.e. To be confident in the results being measured Calibration quantifies and controls errors or uncertainties within measurement processes to an acceptable level therefore there is an ongoing need to maintain the calibration of equipment throughout its lifetime for reliable, accurate and repeatable measurements.

- CALIBRATION

The calibration of equipment shall span to these equipment/apparatus listed below following the standard operating procedure:

- Glass/plastic -wares (standard flasks, pipette, burette, measuring cylinder, syringes)

List of Major Quality Control/Quality Assurance Equipment:

| Equipment | Make | Model | Identification code |
| --- | --- | --- | --- |
| UV Spectrophotometer | Searchtech Instrument , UK | 752N |  |
| Moisture Content Analyser | Ohaus, China | MB120 | BML/QAE/SYP/002 |
| Bucket Autoclave | Optec Lab., China | WT280A | BML/QAE/SYP/003 |
| Analytical weighing balance | S. Mettler, UK | 18011003 | BML/QAE/SYP/004 |
| Viscometer | Viscolead, China | NDJ-5S | BML/QAE/SYP/005 |
| Digital pH meter | Hanna Instruments, China | PHS-3C | BML/QAE/SYP/006 |
| Hand Held pH Meter | Hanna Instruments, China | 211 | BML/QAE/SYP/007 |
| Digital Conductivity meter | Jenway, England | DDS -307 | BML/QAE/SYP/008 |
| Hand Held Conductivity Meter | Exact Instrument , China | CT-3033 | BML/QAE/SYP/009 |
| Heating mantle | Searchtech Instrument , UK | VC006 | BML/QAE/SYP/010 |
| Melting Point Apparatus | Beamex, UK | TC303 | BML/QAE/SYP/011 |
| Water bath | Bright Medicals, England | XMTD-8222 | BML/QAE/SYP/012 |
| Fume cupboard | LABTRAD LABTRADE, Nigeria | BIDI | BML/QAE/SYP/013 |
| Photoelectrical colorimeter | Labtech ,China | LT-01 | BML/QAE/SYP/014 |
| # Magnetic Stirrer Hot Plate | Labtech ,China | 78hw-1 05 | BML/QAE/SYP/015 |
| # Drying Oven 1 | N Medfield Equipment England | DHG-9030A | BML/QAE/SYP/016 |
| # Drying Oven 2 | Medfield Equipment England | DHG-9030A | BML/QAE/SYP/017 |
| Incubator 1 | N Medfield Equipment England | DNP-9052-1 | BML/QAE/SYP/018 |
| Incubatorii Incubator 2 | Medfield Equipment England | DNP-9022-1A | BML/QAE/SYP/019 |
| LAMINAL A Lamina Air Flow Machine | Medifiel Medfield Equipment | SCW-CJ-IF | BML/QAE/SYP/020 |
| Refrigerator 1 | Hisense , Nigeria | 160PR | BML/QAE/SYP/021 |
| Refrigerator 2 | Hisense , Nigeria | 160PR | BML/QAE/SYP/022 |
| Binocular Microscope | Optical Japan | CHC | BML/QAE/SYP/023 |
| Filter integrity Test Apparatus | Houshiji, China | D.O CHROM | BML/QAE/SYP/024 |
| Colony counter | Ohaus, China | QR-005 | BML/QAE/SYP/025 |
| # Electronic weighing balance {500g} | # St. Mettler, UK | # HS-502N | BML/QAE/SYP/026 |
| # Raw material sampler | # Sedol Nig. Ltd, Nigeria | # NA | BML/QAE/SYP/027 |
| # HPLC {High performance Liquid Chromatography | # ECOM Instrument | # 200DH | BML/QAE/SYP/028 |
| # Stability chamber | # RISHABH, India | # RISHmulti20 | BML/QAE/SYP/029 |
| # Sampling booth | # Aeromech, India | # RLAF 1203 SS | BML/QAE/SYP/030 |
| Production Equipment |  |  |  |
| # Dispensing booth | # Aeromech, India | BML/PRE/SYP/001 |  |
| Analytical Weighing balance 200gm | China | BML/PRE/SYP/002 |  |
| Weighing Balance 150kg | China | BML/PRE/SYP/003 |  |
| # Automatic Filling Machine | GMP, India | BML/PRE/SYP/012 |  |
| # Bottle washing machine | GMP, India | BML/PRE/SYP/015 |  |
| Water treatment Equipment |  |  |  |
| UV Water Sterilizer | # Global sterling, India | # BML/WEQ/SYP/007 |  |

List of Major Quality Control/Quality Assurance Equipment:

| Equipment | Make | Model | Identification code |
| --- | --- | --- | --- |
| UV Spectrophotometer | Equiptronics India | 752N | BL/EQ/QAE/001 |
| Auto Polarimeter 1 | Equiptronics India | EQ 805 | BL/ EQ /QAE/002 |
| Analogue Polarimeter 2 | Equiptronics India | EQ 801 | BL/ EQ/QAE/003 |
| Flame Photometer | Equiptronics India | EQ 855A | BL/ EQ /QAE/004 |
| ANALYTIC Analytical weighing balance | Ohaus Ohaus Corporation, USA | AR1530 | BL/ EQ /QAE/005 |
| Digital pH meter | Hanna Instruments Italy | 211 | BL/ EQ /QAE/007 |
| Digital Conductivity meter | Jenway, England | DDS -307 | BL/ EQ /QAE/008 |
| Hand Held Conductivity Meter | Exact Instrument, China | CT -3033 | BL/ EQ /QAE/009 |
| Hand Held pH Meter | Hanna intrumemts,China | PH211 | BL/ EQ /QAE/010 |
| Melting Point Apparatus | Beamex,UK | TC303 | BL/ EQ /QAE/011 |
| Water bath | Surgi Medicals, England | XMTD-8222 | BL/ EQ /QAE/012 |
| Fume cupboard | Labtrade ,Nigeria | BIDI | BL/ EQ /QAE/013 |
| Air borne Particle Counter | Hauda Instrument and Equipment, China | BCJ-1 | BL/ EQ /QAE/014 |
| Magnetic Stirrer Hot Plate | Stuant Scientific , UK | SWAT 05 | BL/ EQ /QAE/015 |
| Hot plate | N Konga, Nigeria | QCP/200A | BL/ EQ /QAE/016 |
| Oven 1 | Vismaram U.K | V1902/STWA6A | BL/ EQ /QAE/017 |
| Oven 2 | Source U.K Kotterman U.K | 2711 | BL/ EQ /QAE/018 |
| Incubator 1 | Eschmed Medical,, England | DNP-9052A | BL/ EQ /QAE/019 |
| Incubatorii Incubator 2 | Gyftyfduyt Gulfex Medical and scientific , England | DNP-9082 | BL/ EQ /QAE/020 |
| Incubator 3 | Medfield Equipment, England | DNP-9052-1 | BL/ EQ /QAE/021 |
| Incubator 4 | Kotterman Kotterman Incubator, Germany | 2735 | BL/ EQ /QAE/022 |
| Potable bucket autoclave 1 | Optec Lab., China | XYXQ280 S.D | BL/ EQ /QAE/023 |
| Potable bucket autoclave 2 | Optec Lab. Optec Lab., China |  | BL/ EQ /QAE/024 |
| Llllllllllllllllll Lamina Air Flow Machine | Medfield E Medfield Equipment, England | SCW-CJ-IF | BL/ EQ /QAE/026 |
| Laminar Air hood | Honri Air clean Tech, China |  | BL/ EQ /QAE/027 |
| Refrigerator | Haier thermocool, Nigeria | HR-142W | BL/ EQ /QAE/028 |
| Filter Integrity Test Apparatus | Houshiji, China | D.O CHROM | BL/ EQ /QAE/031 |
| Digital Scale (500g) | Exact Instrument, China | HS-502N | BL/ EQ /QAE/032 |
| Heating Block | Houshiji, China | TAL-M2 | BL/ EQ /QAE/033 |
| PRODUCTION EQUIPMENT |  |  |  |
| Semi-Automatic Filling Machine 1 | China |  | BL/FR/PE/001 |
| Semi-Auto Semi-Automatic Filling Machine 2 | China |  | BL/FR/PE/002 |
| Semi-Auto Semi-Automatic Filling Machine 3 | China |  | BL/FR/PE/003 |
| Semi-Auto Semi-Automatic Filling Machine 4 | China |  | BL/FR/PE/004 |
| Semi-Auto Semi-Automatic Filling Machine 5 | China |  | BL/FR/PE/005 |
| Digital Weighing Scale 15kg | China |  | BL/SPR/PE/001 |
| Weighing Balance 50kg | China |  | BL/SPR/PE/002 |
| Weighing Weighing Balance 150kg | China |  | BL/SPR/PE/003 |

| EQUIPMENT NAME | GLASSWARE |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

- EQUIPMENT CALIBRATION:

CALIBRATION PARAMETER:

VOLUMETRIC CALIBRATION

| BURETTES |  |  |  |
| --- | --- | --- | --- |
| Nominal capacity, mL |  | 50 | 50 |
| Subdivision, mL |  | 0.1 | 0.1 |
| Tolerance ± mL | CLASS A | 0.05 | 0.05 |
|  | CLASS B | 0.1 | 0.1 |
| MAKE |  |  |  |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |
| Standard deviation |  |  |  |

Status:

| ONE MARK/TRANSFER PIPETTE |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Nominal capacity, mL |  | 25 | 25 | 25 | 25 |
| Tolerance ± mL | CLASS A | 0.03 | 0.03 | 0.03 | 0.03 |
|  | CLASS B | 0.06 | 0.06 | 0.06 | 0.06 |
| MAKE |  |  |  |  |  |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |
| 5 |  |  |  |  |  |
| Standard deviation |  |  |  |  |  |

Status:

| GRADUATED PIPETTE |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Nominal capacity, mL |  | 1 | 5 | 5 | 5 | 10 |
| Subdivision, mL |  | 0.01 | 0.05 | 0.05 | 0.05 | 0.1 |
| Tolerance ± mL | CLASS A | 0.006 | 0.03 | 0.03 | 0.03 | 0.05 |
|  | CLASS B | 0.02 | 0.1 | 0.1 | 0.1 | 0.15 |
| MAKE |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| Standard deviation |  |  |  |  |  |  |

Status:

| VOLUMETRIC FLASK |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Nominal capacity, mL |  | 50 | 100 | 100 | 100 | 200 | 250 | 500 | 500 | 1000 |
| Tolerance ± mL | CLASS A | 0.04 | 0.06 | 0.06 | 0.06 | 0.08 | 0.1 | 0.15 | 0.15 | 0.2 |
|  | CLASS B | 0.08 | 0.15 | 0.15 | 0.15 | 0.18 | 0.2 | 0.3 | 0.3 | 0.8 |
| MAKE |  |  |  |  |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |
| Standard deviation |  |  |  |  |  |  |  |  |  |  |

Status:

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | ANALYTICAL WEIGHING BALANCE |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER : |  |  |  |

CALIBRATION PARAMETER:

WEIGHT CALIBRATION

PRECISION

| s/no | Standard value (g) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 200g | ±0.1% |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |

ACCURACY

| s/no | Date | Standard value (g) | Tolerance | Actual value | Deviation | Remark |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  | 200g | ±0.1% |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  | 20g |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |
| 11 |  | 1g |  |  |  |  |
| 12 |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | ANALYTICAL WEIGHING BALANCE |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

WEIGHT CALIBRATION

PRECISION

| s/no | Standard value (g) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 200g | ±0.1% |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |

ACCURACY

| s/no | Date | Standard value (g) | Tolerance | Actual value | Deviation | Remark |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  | 200g | ±0.1% |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  | 20g |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |
| 11 |  | 1g |  |  |  |  |
| 12 |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | ANALYTICAL WEIGHING BALANCE |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

WEIGHT CALIBRATION

PRECISION

| s/no | Standard value (g) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 200g | ±0.1% |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |

ACCURACY

| s/no | Date | Standard value (g) | Tolerance | Actual value | Deviation | Remark |
| --- | --- | --- | --- | --- | --- | --- |
| 1 |  | 200g | ±0.1% |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  | 20g |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |
| 10 |  |  |  |  |  |  |
| 11 |  | 1g |  |  |  |  |
| 12 |  |  |  |  |  |  |
| 13 |  |  |  |  |  |  |
| 14 |  |  |  |  |  |  |
| 15 |  |  |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | pH METER |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER

MEASUREMENT CALIBRATION

| Buffer 4.01: |  |  |
| --- | --- | --- |
| Reading | Tolerance | Measurement |
| 1 | ±0.10 |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| 6 |  |  |
| 7 |  |  |
| 8 |  |  |
| Standard deviation: |  |  |
| Equipment efficiency (%): NLT 75% : |  |  |

| Buffer 7.01: |  |  |
| --- | --- | --- |
| Reading | Tolerance | Measurement |
| 1 | ±0.10 |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| 6 |  |  |
| 7 |  |  |
| 8 |  |  |
| Standard deviation: |  |  |
| Equipment efficiency (%): NLT 75% : |  |  |

| Buffer 9.18: |  |  |
| --- | --- | --- |
| Reading | Tolerance | Measurement |
| 1 | ±0.10 |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| 6 |  |  |
| 7 |  |  |
| 8 |  |  |
| Standard deviation: |  |  |
| Equipment efficiency (%): NLT 75% : |  |  |

Average efficiency:

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | CONDUCTIVITY METER |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER

MEASUREMENT CALIBRATION

| Cell constant |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Reading | KCl Conductivity µS/cm |  |  | Actual | Remark |
|  | 1413 ± 10, 0.01M |  | 146 ± 5, 0.001M |  |  |
| 1 |  |  |  |  |  |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |
| 5 |  |  |  |  |  |
| 6 |  |  |  |  |  |
| 7 |  |  |  |  |  |
| 8 |  |  |  |  |  |
| 9 |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | UV/VISIBLE SPECTROPHOTOMETER |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

CONTROL OF ABSORBANCE

- Weigh and transfer accurately a quantity not less than 57.0mg and not more than 63mg (60mg) to 1000mL volumetric flask.

- Dissolve and dilute in sufficient 0.005M H2SO4 to produce 1000mL.

Measure the absorbance of the dichromate solution at wavelengths given below, and then calculate the A1%1cm for each wavelength. 
A1%1cm=A*10000weight of potassium dichromate=

| Wavelength | Absorbance Specification | Result | A1%1cm |
| --- | --- | --- | --- |
| 235nm (minima/valley) | 0.748 (0.740 – 0.756) |  |  |
| 257nm (maxima) | 0.865 (0.856 – 0.894) |  |  |
| 313nm (minima/valley) | 0.292 (0.289 – 0.295) |  |  |
| 350 (maxima) | 0.640 (0.634 – 0.646) |  |  |

CALIBRATION OF STRAY LIGHT

- Prepare a 1.2g of potassium chloride previously dried at 1050C for two hours into 100ml volumetric flask

- Add water to dissolve and make up to volume with water.

- Measure the absorbance of the solution at 198.0, 199.0, 200.0, 201.0, 202.0nm using water as blank.

- Acceptance criteria: absorbance greater than 2.00

| S/No | 198.0 | 199.0 | 200.0 | 201.0 | 202.0 |
| --- | --- | --- | --- | --- | --- |
| Set 1 |  |  |  |  |  |
| Set 2 |  |  |  |  |  |
| Set 3 |  |  |  |  |  |

CALIBRATION OF RESONANCE POWER

- Prepare a 0.02% v/v toluene in hexane.

- Scan through wavelength 260nm to 420nm.

- Acceptance criteria: the ratio of the absorbance at maximum at about 269nm to that at the minimum at about 266nm is not less than 1.5.

| Parameter | run 1 | run 2 | run 3 | mean |
| --- | --- | --- | --- | --- |
| Absorbance at 269nm |  |  |  |  |
| Absorbance at 266nm |  |  |  |  |
| ratio |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | FUME CUPBOARD |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

| SASH OBSTRUCTION: |  |
| --- | --- |
| Is the sash sufficiently transparent, not etched from acid or marked by any acid? | Comment |
| MAXIMUM APERTURE HEIGHT |  |
| Maximum aperture height between the base of the cupboard and the underside of the sash, NMT 0.5m (maximum working aperture marked) | result |
| FACE VELOCITY TESTING: |  |
| Using an anemometer, air flow velocity (0.4 – 1.0m/s), maximum deviation ± 20% from mean within the grid | result |
| AIR FLOW INDICATOR: |  |
| Air flow indicator of any sort? | observation |
| LUMINOSITY: |  |
| NLT 500lux |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | WATER BATH |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

TEMPERATURE CALIBRATION

Set the water-bath to 37oC, 50 oC and 90 oC and check against a standard thermometer.

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 37 | ±0.5 oC |  |  |
| 2 | 50 | ±0.2 oC |  |  |
| 3 | 90 | ±0.1 oC |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | MAGNETIC STIRRER |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

TEMPERATURE CALIBRATION:

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 10 | ±2oC |  |  |
| 2 | 20 | ±2oC |  |  |
| 3 | 30 | ±2oC |  |  |
| 4 | 40 | ±2oC |  |  |
| 5 | 50 | ±2oC |  |  |
| 6 | 60 | ±2oC |  |  |
| 7 | 70 | ±2oC |  |  |
| 8 | 80 | ±2oC |  |  |
| 9 | 90 | ±2oC |  |  |
| 10 | 100 | ±2oC |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | MELTING POINT APPARATUS |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

MEASUREMENT CALIBRATION:

While using the instrument’s SOP, determine the melting point of these reference materials.

| s/no | Reference Substance | Standard range | Measurement |
| --- | --- | --- | --- |
| 1 | Ibuprofen | 75 – 78oC |  |
| 2 | Nicotinamide | 128 - 131 oC |  |
| 3 | Sulfamethoxazole | 169 - 172 oC |  |
| 4 | Trimethoprim | 199 - 203 oC |  |
| 5 | Diclofenac Sodium | 277 – 283oC |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | PHOTOELECTRIC COLORIMETER |
| --- | --- |
| EQUIPMENT CODE |  |
| INSTRUMENT MAKE |  |
| SOPs NUMBER: |  |

CALIBRATION PARAMETER

LINEARITY CHECK:

- Weigh and transfer accurately a quantity not less than 57.0mg and not more than 63mg (60mg) to 1000mL volumetric flask.

- Dissolve and dilute in sufficient 0.005M H2SO4 to produce 1000mL. Determine the maximum wavelength of the solution while zeroing with 0.005M H2SO4.

- Using this wavelength, prepare five dilution of this solution, determine the optical density of these solutions. (5mL, 10mL, 15mL, 20mL, 25mL)

| Maximum wavelength: |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| s/no | Concentration mg/mL | Optical density |  |  | Mean |
| 1 | 0.3 |  |  |  |  |
| 2 | 0.6 |  |  |  |  |
| 3 | 0.9 |  |  |  |  |
| 4 | 1.2 |  |  |  |  |
| 5 | 1.5 |  |  |  |  |
| Linearity equation |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | OVEN |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

Set the water-bath to 50oC, 105 oC and 150 oC and check against a standard thermometer.

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 50 | ±2% |  |  |
| 2 | 105 | ±2% |  |  |
| 3 | 150 | ±2% |  |  |
| 4 | Alarm challenge testing |  |  |  |

CONCLUSION/OBSERVATION

| EQUIPMENT NAME | OVEN |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

Set the water-bath to 50oC, 105 oC and 150 oC and check against a standard thermometer.

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 50 | ±2% |  |  |
| 2 | 105 | ±2% |  |  |
| 3 | 150 | ±2% |  |  |
| 4 | Alarm challenge testing |  |  |  |

CONCLUSION/OBSERVATION

| EQUIPMENT NAME | MOISTURE CONTENT ANALYZER |  |  |
| --- | --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION |  |
| INSTRUMENT MAKE |  |  |  |
| SOPs NUMBER: |  |  |  |

CALIBRATION PARAMETER:

TEMPERATURE CALIBRATION

Operate the instrument as per SOP and calibrate the temperature to 60oC, 105oC and 150oC and check against a standard thermometer.

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 60 | ±2oC |  |  |
| 2 | 105 | ±2oC |  |  |
| 3 | 150 | ±2oC |  |  |

WEIGHT CALIBRATION

Operate the instrument as per SOP and calibrate the weight with 0.5g and 1g

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 0.5g | ±0.1% |  |  |
| 2 |  |  |  |  |
| 3 |  |  |  |  |
| 4 |  |  |  |  |
| 5 |  |  |  |  |
| 6 |  |  |  |  |
| 7 |  |  |  |  |
| 8 |  |  |  |  |
| 9 |  |  |  |  |
| STANDARD DEVIATION |  |  |  |  |

| s/no | Standard value (oC) | Tolerance | Actual value | Remark |
| --- | --- | --- | --- | --- |
| 1 | 1g | ±0.1% |  |  |
| 2 | 1g | ±0.1% |  |  |
| 3 | 1g | ±0.1% |  |  |
| STANDARD DEVIATION |  |  |  |  |

MEASUREMENT CALIBRATION

| s/no | WEIGHT OF WATER | LIMIT | MOISTURE CONTENT (%) |
| --- | --- | --- | --- |
| 1 |  | 100 ± 2% |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |
| 6 |  |  |  |
| 7 |  |  |  |
| 8 |  |  |  |
| 9 |  |  |  |
| STANDARD DEVIATION |  |  |  |
| %RELATIVE STANDARD DEVIATION |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | REFRIGERATOR |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER:

TEMPERATURE CALIBRATION:

Using the temperature data logger, suspend the sensor in the freezer/refrigerator for three hours.

| Ambient temperature | 0min | 30 min |  | 60 min | 90 min |  | 120 min |  | 150 min | 180 min |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Freezer setting | Tolerance |  | actual |  |  |  |  | remark |  |  |
|  |  |  | freezer |  |  | Fridge |  |  |  |  |
| 0 | ±2OC |  |  |  |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | REFRIGERATOR |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER:

TEMPERATURE CALIBRATION:

Using the temperature data logger, suspend the sensor in the freezer/refrigerator for three hours.

| Ambient temperature | 0min | 30 min |  | 60 min | 90 min |  | 120 min |  | 150 min | 180 min |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Freezer setting | Tolerance |  | actual |  |  |  |  | remark |  |  |
|  |  |  | Freezer |  |  | Fridge |  |  |  |  |
| 0 | ±2OC |  |  |  |  |  |  |  |  |  |
| 1 |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |  |  |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | STABILITY CHAMBER |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER

TEMPERATURE AND HUMIDITY PERFORMANCE CALIBRATION:

Using the temperature and humidity sensor data logger, suspend the sensor in the chamber. Set the temperature and humidity accordingly while obeying operational qualifications.

| Ambient temperature |  |  |  | Ambient humidity |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set oC | Set %RH |  | Tolerance | Actual oC | Actual %RH | Deviation |  |  | Remark |
| 30 | 50 |  | ±1.5oC ±3.0%RH |  |  |  |  |  |  |
| 45 | 50 |  |  |  |  |  |  |  |  |
| 55 | 50 |  |  |  |  |  |  |  |  |
| 30 | 65 |  |  |  |  |  |  |  |  |
| 45 | 65 |  |  |  |  |  |  |  |  |
| 55 | 65 |  |  |  |  |  |  |  |  |
| 30 | 75 |  |  |  |  |  |  |  |  |
| 45 | 75 |  |  |  |  |  |  |  |  |
| 55 | 75 |  |  |  |  |  |  |  |  |

ALARM CHALLENGE TESTING:

| s/no | Alarm challenge | OK | Not OK |
| --- | --- | --- | --- |
| 1 | Temperature high |  |  |
| 2 | Temperature low |  |  |
| 3 | Humidity high |  |  |
| 4 | Humidity low |  |  |
| 5 | Door open |  |  |
| 6 | Low water level |  |  |
| LUMINOSITY |  |  |  |
| 1 | Tray 1 | NLT 50lux |  |
| 2 | Tray 2 | NLT 50lux |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | HPLC |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER

LEAKAGE TEST BY PRESSURE DROP (PUMP CALIBRATION)

| Maximum pressure | Set pressure | Start time | Stop time | Total time | Observed pressure | remark |
| --- | --- | --- | --- | --- | --- | --- |
| 300bar |  |  |  |  |  |  |

Status: accepted/declined

FLOW RATE CALIBRATION (PUMP CALIBRATION)

| Flow rate (ml/min) | Time required to collect 10mL mobile phase |  |  | Limit | Remark |
| --- | --- | --- | --- | --- | --- |
|  | Theoretical (s) | Water Actual | Methanol actual |  |  |
| 0.5 | 1200 |  |  | 1194-1206 |  |
| 1.0 | 600 |  |  | 594-606 |  |
| 1.5 | 450 |  |  | 443-457 |  |
| 2.0 | 300 |  |  | 294-306 |  |

Status: accepted/declined

REPRODUCIBILITY, LINEARITY OF INJECTION VOLUME AND LINEARITY OF DETECTOR RESPONSE

| Solution | Volume of benzene (mL) | Volume of toluene (mL) | Diluted to with methanol to mL |
| --- | --- | --- | --- |
| 1 | 1.0 | 0 | 50 |
| 2 | 0 | 1.0 | 50 |
| 3 | 0.5 | 1.0 | 50 |
| 4 | 1.0 | 1.0 | 50 |
| 5 | 1.5 | 1.0 | 50 |

| Parameters | Test condition | Applied condition |
| --- | --- | --- |
| Mobile phase | Methanol: water (70:30) |  |
| HPLC column | C18 |  |
| Flow rate | 1.0mL/min |  |
| wavelength | 254nm |  |

| Injection volume | Parameter | Injection 1 |  | Injection 2 |  | Injection 3 |  | Mean |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Benzene | Toluene | Benzene | Toluene | Benzene | Toluene |  |
| 10 | RT |  | NA |  | NA |  | NA |  |
|  | Area |  | NA |  | NA |  | NA |  |
| 10 | RT | NA |  | NA |  | NA |  |  |
|  | Area | NA |  | NA |  | NA |  |  |
| 10 (3) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 10 (4) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 10 (5) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 15 (3) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 15 (4) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 15 (5) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 20 (3) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 20 (4) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 20 (5) | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 25 | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| 30 | RT |  |  |  |  |  |  |  |
|  | Area |  |  |  |  |  |  |  |
| Standard deviation |  |  |  |  |  |  |  |  |
| Linearity equation |  |  |  |  |  |  |  |  |
| Correlation coefficient |  |  |  |  |  |  |  |  |

Acceptance criteria: NLT 0.99

Status:

D2 LAMP ENERGY CHECK (DETECTOR)

| Set wavelength | Limit | Reference energy observed |
| --- | --- | --- |
| 254 | NLT 200 |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | BUCKET AUTOCLAVE |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER;

Calibration of autoclave using the biological indicator (Bacillus stereothermophilus), using FTM medium.

| Strip | Sterilisation condition | Expectation | Result |
| --- | --- | --- | --- |
| 1 | 121oC, 15lbs, 15 minutes | No growth after 14days |  |
| 2 |  |  |  |

Calibration of the autoclave using Heat labile physical indicator:

| Strip | Sterilisation condition | Expectation | Result |
| --- | --- | --- | --- |
| 1 | 121oC, 15lbs, 15 minutes | Within “accept” limit |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |

(Attach a sample of the chemical strip)

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | INCUBATORS |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER

Temperature calibration:

| 20 – 25oC INCUBATOR: 25OC |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| THERMOMETER | LOCATION |  |  |  |  |
|  | Time (min) | A | B | C | D |
| 1 | 30 |  |  |  |  |
| 2 | 60 |  |  |  |  |
| 3 | 90 |  |  |  |  |
| 4 | 120 |  |  |  |  |

Acceptance criteria: 20 – 26oC

Status:

| 35 – 37oC INCUBATOR: 37OC |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| THERMOMETER | LOCATION |  |  |  |  |
|  | Time (min) | A | B | C | D |
| 1 | 30 |  |  |  |  |
| 2 | 60 |  |  |  |  |
| 3 | 90 |  |  |  |  |
| 4 | 120 |  |  |  |  |

Acceptance criteria: 34 – 38oC

Status:

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | BACTERIOLOGICAL COLONY COUNTER |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER

Read-out Calibration:

Using the SOP for the Calibration of the Bacteriological Colony Counter, verify the read-out using the number stencil.

| s/no | Surface pressed | Readout | Button pressed | Readout |
| --- | --- | --- | --- | --- |
| 1 | 100 times |  | 100 times |  |
| 2 | 100 times |  | 100 times |  |
| 3 | 100 times |  | 100 times |  |
| 4 | 100 times |  | 100 times |  |
| 5 | 100 times |  | 100 times |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | DISPENSING BOOTH |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER:

Velocity and Luminosity

| Parameter | Specification |  |
| --- | --- | --- |
| Luminosity | NLT 50lux |  |
| Velocity | NLT 0.4m/s |  |

Airflow laminarity by smoke test

| Smoke chemical |  | Chemical ID |  |
| --- | --- | --- | --- |
| Pressure differential |  |  |  |
| Smoke generated? |  |  |  |
| Laminarity observed when put in the OFCF area from left to right? |  |  |  |
| Repeat the smoke test three more times (upper, middle and lower), compliance? |  |  |  |
| Video evidence? (reference) |  |  |  |

Non-viable particulate count

| Particle size | Cumulative particle concentration per cubic meter |  |  |
| --- | --- | --- | --- |
|  | 0.5µm | 1.0 µm | 5.0 µm |
| ISO Class 5 | 3520 | 832 | 29 |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |

Status:

Viable particle count

| Air sampler: active air sampling |  |  |
| --- | --- | --- |
| Sampling locations |  | TAMC |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| Acceptance Criteria: ≤1cfu per cubic meter |  |  |
| Settle Plate: Passive air sampling. |  |  |
| Sampling locations | TAMC | TYMC |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| Acceptance Criteria: ≤1cfu per plate, 90Ø. |  |  |
| Exposure time: 4 hours. |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | SAMPLING BOOTH |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PARAMETER:

Velocity and Luminosity

| Parameter | Specification |  |
| --- | --- | --- |
| Luminosity | NLT 50lux |  |
| Velocity | NLT 0.4m/s |  |

Airflow laminarity by smoke test

| Smoke chemical |  | Chemical ID |  |
| --- | --- | --- | --- |
| Pressure differential |  |  |  |
| Smoke generated? |  |  |  |
| Laminarity observed when put in the OFCF area from left to right? |  |  |  |
| Repeat the smoke test three more times (upper, middle and lower), compliance? |  |  |  |
| Video evidence? (reference) |  |  |  |

Non-viable particulate count

| Particle size | Cumulative particle concentration per cubic meter |  |  |
| --- | --- | --- | --- |
|  | 0.5µm | 1.0 µm | 5.0 µm |
| ISO Class 5 | 3520 | 832 | 29 |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |
| 5 |  |  |  |

Status:

Viable particle count

| Air sampler: active air sampling |  |  |
| --- | --- | --- |
| Sampling locations |  | TAMC |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| Acceptance Criteria: ≤1cfu per cubic meter |  |  |
| Settle Plate: Passive air sampling. |  |  |
| Sampling locations | TAMC | TYMC |
| 1 |  |  |
| 2 |  |  |
| 3 |  |  |
| 4 |  |  |
| 5 |  |  |
| Acceptance Criteria: ≤1cfu per plate, 90Ø. |  |  |
| Exposure time: 4 hours. |  |  |

CONCLUSION/OBSERVATION:

| EQUIPMENT NAME | POLARIMETER ( DIGITAL) |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PROCEDURE

Materials – dextrose monohydrate, distilled water, 5M Ammonia solution

Press clear button to erase all previous stored values.

Fill the sample tube with blank (water) and place it in the compartment.

Press start, wait for the display to give the reading, repeat and press average

Then, press zero, display will read 0000.

3.3g of dextrose monohydrate was weighed and dissolved in 100ml flask

0.2ml of 5M Ammonia solution was added and then made to the mark of the flask with distilled water

The polarimeter tube was rinsed with distilled water and also with the prepared solution.

The tube was filled with the solution in such a way that the bubble is allowed in the tube, placed the tube back to the appropriate side, then press start

The reading was taken for solution in the polarimeter

The above procedure was taken for 5.5g, 11g, 22g, and 55g (Dextrose Monohydrate)

5.0g, 10g, 20g, and 50g (Dextrose Anhydrous)

Results

| s/n | Concentration (% w/v) |  | Results |
| --- | --- | --- | --- |
| 1 | Anhydrous | Monohydrate |  |
|  | 3 | 3.3 |  |
| 2 | 5 | 5.5 |  |
| 3 | 10 | 11 |  |
| 4 | 20 | 22 |  |
| 5 | 50 | 55 |  |

| EQUIPMENT NAME | POLARIMETER ( ANALOQUE) |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PROCEDURE

Materials – AR Sucrose/dextrose monohydrate, distilled water, 5M Ammonia solution

Polaroid wheel was rotated so that display reads Zero

The tube was filled with distilled water and placed it in the chamber

The wheel was rotated back and forth so that intensity of light is minimum.

At this position, the display should read zero, repeat step 2 and3 repeatedly

3.3g of dextrose monohydrate was weighed and dissolved in 100ml flask

0.2ml of 5M Ammonia solution was added and then made to the mark of the flask with distilled water

The polarimeter tube was rinsed with distilled water and also with the prepared solution.

The tube was filled with the solution in such a way that the bubble is allowed in the tube, placed the tube back to the appropriate side

The reading was taken for solution in the polarimeter

The above procedure was taken for 5.5g, 11g, 22g, and 55g (Dextrose Monohydrate)

5.0g, 10g, 20g, and 50g (Dextrose Anhydrous)

Results

| s/n | Concentration (% w/v) |  | Results |
| --- | --- | --- | --- |
| 1 | Anhydrous | monohydrate |  |
|  | 3 | 3.3 |  |
| 2 | 5 | 5.5 |  |
| 3 | 10 | 11 |  |
| 4 | 20 | 22 |  |
| 5 | 50 | 55 |  |

| EQUIPMENT NAME | FLAME PHOTOMETER |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

CALIBRATION PROCEDURE

Materials – Sodium Chloride, Potassium Chloride, distilled water

The mains was on and the nebulizer were dipped inside the distilled water

The switch was on and the fuel knob

Clean, pulse free air was then set at a pressure between 4-6psi

The fuel line was put on and the lighter was on

The fuel line was regulated until blue flame was achieved.

Allow the flame to stabilize for 1-2minutes

The instrument was allowed to warm up for 20 mins

Various concentration of the standard solution of potassium and sodium ions were prepared by dissolving appropriate quantity in deionised water.

The concentration of 2.0ppm, 4.0ppm, 6.0ppm, 8.0ppm and 10.0ppm was prepared

For 10.0ppm concentration, the digital was set to 000 display using the fine and coarse sensitivity knob.

The reading for other concentration were then taken ( 2ppm on ward)

The graph of concentration against digital display was then plotted using the readings obtained.

Results

| s/n | Concentration (ppm) | Digital reading for Na+ | Digital reading for K+ |
| --- | --- | --- | --- |
| 1 | 2.0 |  |  |
| 2 | 4.0 |  |  |
| 3 | 6.0 |  |  |
| 4 | 8.0 |  |  |
| 5 | 10.0 |  |  |

| EQUIPMENT NAME | FILLING MACHINE |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

PARAMETER CALIBRATION

VOLUMETRIC CALIBRATION

| Volume | 100mL |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Nozzle head | 1 | 2 | 3 | 4 | 5 | 6 |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |

Average of the last 5 run: 100 ± 5mL

| Volume | 200mL |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Nozzle head | 1 | 2 | 3 | 4 | 5 | 6 |
| 1 |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |
| 4 |  |  |  |  |  |  |
| 5 |  |  |  |  |  |  |
| 6 |  |  |  |  |  |  |
| 7 |  |  |  |  |  |  |
| 8 |  |  |  |  |  |  |
| 9 |  |  |  |  |  |  |

Average of the last 5 run: 200 ± 10Ml

| EQUIPMENT NAME | BOTTLE WASHING MACHINE |  |
| --- | --- | --- |
| EQUIPMENT CODE |  | LOCATION: |
| INSTRUMENT MAKE |  |  |
| SOPs NUMBER: |  |  |

| BOTTLE WASHING @400C |  |  | BOTTLE WASHING @600C |  |  |
| --- | --- | --- | --- | --- | --- |
| MICROBIAL LIMITS |  | RESULT | MICROBIAL LIMITS |  | RESULT |
| Bacteria | NMT |  | Bacteria |  |  |
| Mould/yeast |  |  | Mould/yeast |  |  |
| Coliform |  |  | Coliform |  |  |

| AREA | Production area |
| --- | --- |
| LOCATION: |  |

CONCLUSION/OBSERVATION:

AREA INSPECTION:

Dispensing Area:

| ENVIRONMENTAL CONDITION |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TEMPERATURE (OC) |  |  |  |  |  |  |  |
| RELATIVE HUMIDITY (%) |  |  |  |  |  |  |  |
| LUMINOSITY |  |  |  |  |  |  |  |
| PARTICLE COUNT (min. ISO CLASS 8) |  | 0.5 |  | 1.0 |  | 5.0 |  |
|  |  | 3520000 |  | 832000 |  | 29300 |  |
| Bio-burden | Limit | Settle plates (diameter 90mm; NMT 100 cfu/ 4 hours) Air sample (NMT 200cfu/m3) |  |  |  |  |  |
|  | Active |  |  |  |  |  |  |
|  | Passive |  |  |  |  |  |  |

Mixing Area:

| ENVIRONMENTAL CONDITION |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TEMPERATURE (OC) |  |  |  |  |  |  |  |
| RELATIVE HUMIDITY (%) |  |  |  |  |  |  |  |
| LUMINOSITY |  |  |  |  |  |  |  |
| PARTICLE COUNT (min. ISO CLASS 8) |  | 0.5 |  | 1.0 |  | 5.0 |  |
|  |  | 3520000 |  | 832000 |  | 29300 |  |
| Bio-burden | Limit | Settle plates (diameter 90mm; NMT 100 cfu/ 4 hours) Air sample (NMT 200cfu/m3) |  |  |  |  |  |
|  | Active |  |  |  |  |  |  |
|  | Passive |  |  |  |  |  |  |

Filling area:

| ENVIRONMENTAL CONDITION |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TEMPERATURE (OC) |  |  |  |  |  |  |  |
| RELATIVE HUMIDITY (%) |  |  |  |  |  |  |  |
| LUMINOSITY |  |  |  |  |  |  |  |
| PARTICLE COUNT (min. ISO CLASS 8) |  | 0.5 |  | 1.0 |  | 5.0 |  |
|  |  | 3520000 |  | 832000 |  | 29300 |  |
| Bio-burden | Limit | Settle plates (diameter 90mm; NMT 100 cfu/ 4 hours) Air sample (NMT 200cfu/m3) |  |  |  |  |  |
|  | Active |  |  |  |  |  |  |
|  | Passive |  |  |  |  |  |  |

CONCLUSION/AREA:

7.0 DEVIATION:

8.0 SUMMARY/CONCLUSION:

9.0 REPORT ENDORSEMENT:

| LABORATORY HEAD | Name: | Signature/date: |
| --- | --- | --- |
| QUALITY ASSURANCE MANAGER | Name: | Signature/date: |