- OBJECTIVE:

To provide written guidelines for cleaning and operation of Air Particulate Counter

- SCOPE:

This SOP covers the process of cleaning and Operation of Air Particulate Counter

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled

4.0 ABBREVIATIONS AND DEFINITIONS

- LCD – Liquid Crystal Display

- PROCEDURE

Cleaning:

To self-clean

- Connect sample air inlet to self-cleaning port with sample tube

- Press the measure key to clear all the size channels counts to zero indicating no false count of the instrument exist

Operation:

- Connect the switch to main electrical supply

- Switch on the power by pressing the power switch

- Press configuration key, then pressing Adjust key to correct time and date

- Pre heat for 5 minutes

- Select the measure key for the air pump to work, the flow rate and the measuring data are displayed on the LCD with total counting mode

- Connect sample air inlet to a sample probe

- Adjust the flow rate to 2.83L/min by rotating flow control knob

- Press measure key to stop measuring and reset key to re-start a new sample cycle

- Ensure that the area is cleaned, and air handling unit is kept ON for at least two hours before taking particle count

- The locations for taking the samples shall be near doors, return air raisers, at corners and near equipment

- Disconnect the cable from the main electricity supply

- The particle count testing shall be carried out as per schedule

- The cumulative particle concentration for the area should not exceed the desired grade limit as shown in the table below.

| Grade | Maximum concentration limits(particulate/m3) for particle> sizes zone |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | At rest |  |  |  | In operation |  |  |  |
|  | ≥0.5µm |  | ≥ 5.0µm |  | ≥0.5µm |  | ≥ 5.0µm |  |
|  | Spec | Reading | spec | Reading | Spec | Reading | spec | Reading |
| A | 3500 |  | 0 |  | 3500 |  | 0 |  |
| B | 3500 |  | 0 |  | 350000 |  | 2000 |  |
| C | 350000 |  | 2000 |  | 3500000 |  | 20000 |  |
| D | 3500000 |  | 20000 |  | Not defined |  | Not defined |  |

- Record the usage activities in the user logbook LGB/QDG/MC/014/01 for the Air Particulate counter appropriately.

Air Particulate Counter User Logbook LGB/QDG/MC/014/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

- REFERENCE/ENCLOSURE

- Air Particulate Counter Equipment user manual

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |