- OBJECTIVE:

To provide written guidelines for Cleaning and Operation of Active Air Sampler.

- SCOPE:

This SOP covers the process of Cleaning and Operation of Active Air Sampler.

3.0 RESPONSIBILITY/ACCOUNTABILITY:

- The microbiologist must ensure that the equipment is kept clean

- The user reports any equipment malfunction.

- The Quality control Manager ensures external calibration is carried out as scheduled.

4.0 ABBREVIATIONS AND DEFINITIONS

- TSA - Tryptic soya agar

- CFU - Colony forming unit

- PROCEDURE

Cleaning

- Switch ‘OFF’ the Active air sampler

- Clean with 70% IPA

Operation

- Ensure that the Active air sampler is clean.

- The sanitized stainless-steel hood and sterilized petri plates (containing solidified media) are kept in a clean box as per required.

- The active air sampler and the plates are carried into the sampling area.

- The pre-incubated sterilized labeled plate is being placed in the feeder cone circular clamp assembly in position.

- The air sampler is then placed in the area where air sampling is being done.

- The ‘ON’ switch is then turned on which is provided on the air sampler.

- Initial display shall be observed.

- Press ‘START ON‘  the control panel to obtain 100 liter air.

- The instrument shall be automatically   OFF after sucking 100 liter of air.

- Follow the guide below for ALERT/ ACTION LEVELS

| GRADE LEVEL | ALERT LEVEL (cfu/m3) | ACTION LEVEL (cfu/m3) |
| --- | --- | --- |
| A | N/A | ≥1 |
| B | 3-10 | >10 |
| C | 50-99 | >100 |
| D | 100-200 | >200 |

- Record the result in the Infusion active air sampling record form BML-QDG-MC-013-F01 or Syrup active air sampling record form BML-QDG-MC-013-F02 respectively.

- Then open the stainless-steel hood and remove the plate aseptically.

- For next location, sanitize the stainless-steel hood and carry within air sampler trolley.

- Cover the plate with their upper lid then incubate for 48 hours at 30-35oC and 72 hours at 20-25˚C respectively.

- Record the usage activities in the user logbook LGB/QDG/MC/013/01 for the Air sampler machine appropriately.

Air Sampler machine User Logbook LGB/QDG/MC/013/01

| Date | Used for | Start time | End time | Total time use | Status | Operator sign | Checked by | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

PRECAUTIONS

- Switch OFF when the instrument in not in use.

Materials, Equipment, Apparatus and Reagents

- PPE’s

- Active air sampler

- Sterilized media

- 70% IPA

- Cotton wool

- Hand gloves

Safety, Health and Environment Precautions

- Wear appropriate PPE

FREQUENCY - Twice a week

- REFERENCE/ENCLOSURE

- Infusion Active air sampling record form - BML-QDS-MC-013-F01

- Syrup Active air sampling record form - BML-QDS-MC-013-F02

- International journal of trend scientific research and development

- Air Sampler Equipment user manual

- TRAINING INVOLVEMENT

| Quality Control Head | Quality control microbiologist, Laboratory assistant |
| --- | --- |

8.0 DISTRIBUTION AND DOCUMENTATION CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | QUALITY CONTROL LABORATORY |

9.0 REVISION HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| ### 00 | ### Nil | First issue |
| ### 02 | ### 01 | To Update |
| ### 03 | ### 02 | - Inclusion of enclosures format - Change in sop format generally |