- OBJECTIVE:

This SOP establishes a consistent procedure for calibrating Ohaus Adventurer balances, ensuring accurate measurements for laboratory or industrial use.

- SCOPE:

This SOP describes the written guidelines for calibration and operation of Ohaus Adventurer Balance.

- RESPONSIBILITIES/ACCOUNTABILITIES:

QC Analyst: Follows SOP, performs calibration, and records results.

QCH: ensures availability of standard weights, reviews records.

QAM: Approves SOP

- ABBREVIATION AND DEFINITIONS

QC- Quality Control

QCH- Quality Control Head

QAM- Quality Control Manager

- EQUIPMENT AND MATERIAL

- Ohaus Adventurer Balance

- Certified, traceable standard weights (e.g., 100g, 50g, or as required by model)

- Balance cleaning supplies (soft brush, lint-free cloth)

- PROCEDURES

PRECAUTION

- Work in a stable, vibration-free area.

- Allow balance to warm up for 20 -30mins.

- Ensure balance is level (check bubble).

- Use clean weights and pan.

- Avoid touching weights with bare hands.

EQUIPMENT CALIBRATION

  - Linearity calibration should be carried out at least once daily to minimize deviation between actual and displayed weights within the weighing range of the balance.  Calibration should be done before any weighing operation if the substance to be weighed is for analytical purposes or intended for preparation of standard solutions.

  - Clean the balance with soft cloth and take extra care to clean the pans so as to remove accumulated residue of materials previously weighed.

  - Connect the balance to the electricity source

  - With the balance OFF, press and hold >O/T< until MENU is displayed, then release it and UNITS is displayed. Press MODE OFF, LIN is displayed

  - Ensure the balance reads "0.0000g" or shows stable zero.

  - Press >O/T<, -C- is displayed followed by the value of the first mass which must be placed on the pan

Note: Do not disturb the balance while –C- is displayed. Carefully place the specified standard weight (e.g., 100g) onto the center of the pan using tweezers.

  - Place the second required mass on the pan and momentarily press >O/T<-C- is displayed. When the weight on the pan is displayed with the stability indicator, the balance is calibrated and returns to weighing mode

  - Remove the calibration masses from the pan.

OPERATION

  - Press >O/T< to turn the balance on. It is ready to weigh in grams

  - Press the >O/T< momentarily to zero the display if necessary

  - Place item to be weighed on the pan and read the weight on the display. The stability indicator * appears when the reading is stable

  - When weighing items that must be held in a container, tar to subtract the weight of the container from the total weight of the pan. With an empty container on the pan, press >O/T< to zero the display

  - As material is added to the container, the net weight is displayed. The tarred weight remains in the balance memory until >O/T<is pressed again.

  - When the balance is not in use, switch off by pressing and holding MODE OFF.

- REFERENCE AND ENCLOSURE

ENCLOSURE 1- WEIGHING BALANCE CALIBRATION LOG - BML-QDG-QA-033-F01

- TRAINING INVOLVEMENT

| Quality Assurance manager | Quality Control Head |
| --- | --- |
| QA Doc. Unit staff | Quality Control Chemists |

- DISTRIBUTION AND CONTROL:

| MASTER COPIES | QUALITY ASSURANCE |
| --- | --- |
| CONTROLLED COPIES | CHEMICAL LABORATORY |

- CHANGE HISTORY

| CURRENT REVISION | SUPERSEDE | REASON FOR REVIEW |
| --- | --- | --- |
| 00 | Nil | First issue |
| 06 | 05 | - Change in sop format. - Inclusion of traceability of enclosures to parent SOP |

ENCLOSURE 1

Ohaus Adventurer Calibration Log

|  | Biomedical Limited, No. 1 Ohimege Road, Industrial Estate, Ilorin, Kwara State, PMB 1449 | Doc.No.: BML/QDF/QA/033 |
| --- | --- | --- |
|  | QUALITY ASSURANCE DEPARTMENT EQUIPMENT CALIBRATION LOG OHAUS ADVENTURER BALANCE | Version No.: 02 |

Doc. No.; BML/QDG/QA/033/F01

| Date | Standard weight used (g)(kg) | Readings | Deviation | Specification (+ 0.1%) | Remark | Done By | Checked by |
| --- | --- | --- | --- | --- | --- | --- | --- |

_____________________

Q.A Manager

Sign /Date